## Slide 1: To hide, or not to hide, that is the question Towards common philosophy about rejecting, hiding badqual of the calibration data
Reason for rejecting: to assure that only certified/good calibrations are used to be associated with science data (science ABs, the static solution of the calSelector,) and only certified/good master calibrations are ingested in the Archive;
Reason for hiding: to clean the ESO Archive from bad/unusable calibration raw frames that could otherwise be picked up by future users (the dynamic solution of the calSelector);
Reason for badqual: to assure that, on the calChecker level, only certified/good calibrations are considered for science association; main motivation is to alert Paranal if good calibrations are missing;
Good data – definition: stable in time, closely correspond to the status of the instrument;
should be “scored green”, but can also be “scored red” if the reasons are understood
and accepted;

## Slide 2
Data | REJECT | VCAL rm | HIDE | BADQUAL
saturated, severely underexposed, contaminated by artifacts, unusable/bad for science | YES | YES | YES | YES
pipeline failed because data are faulty/bad |  | YES | YES | YES
“scored red” but reflect out of ordinary, acceptable status of the instrument; correspond to the science data | NO | NO | NO | NO
“scored red” because reflect inacceptable, out of ordinary status of the instrument; not needed for science | YES | YES | YES | NO
“scored red” but show only slight/acceptable departure from thresholds; useful for science | NO | NO | NO | NO
“scored red” but can still be useful for science; calibrations cannot be replaced | NO but open ticket to USD | NO | NO | NO
“scored red” but still useful; new, good set was re-taken | YES | YES | NO | NO
“scored red”, but may still be useful (e.g. IR twilight flats sampling too narrow dynamic range) | YES | YES | NO | YES
pipeline failed but data fine, usable for science (e.g. recipe cannot handle instr. setup, incomplete template) |  | YES | NO | YES
pipeline failed, data fine but not usable for science (e.g. incomplete template) |  | YES | YES | YES

## Slide 3
FOOTNOTES:
Hiding of the raw frames in the Archive should be exception and not a habit.
The assumption is that Paranal delivers mainly good data.