ESO
Interface Control Document

between the Electro-Mechanical Hardware and the Control System of the ATS


Doc:

Issue

Date

Page
VLT-ICD-ESO-15100-1528
4.3

15/05/01

76 of 80


[image: image1.wmf]EUROPEAN SOUTHERN OBSERVATORY


Organisation Européenne pour des Recherches Astronomiques dans l'Hémisphère Austral 

Europäische Organisation für astronomische Forschung in der südlichen Hemisphäre
VERY  LARGE  TELESCOPE

Interface Control Document

between the

Electro-Mechanical Hardware and the Control System

of the Auxiliary Telescope System

Doc. No.: VLT-ICD-ESO-15100-1528

Issue: 4.3

Date: 15/05/01


Prepared:  G.Chiozzi, K.Wirenstrand, P.Duhoux, 


R.Karban, J-M.Moresmau, M.Duchateau



Name



Date

Signature


Approved:  B. Koehler, G.Raffi

Name



Date

Signature


Released:  M. Tarenghi

Name



Date

Signature

VLT PROGRAMME * TELEPHONE: (089) 3 20 06-0 * FAX: (089) 3 20 06 514

CHANGE RECORD

ISSUE
DATE
SECTION/PAGE 

AFFECTED
REASON/INITIATION

DOCUMENTS/REMARKS



1
081297
All
Issue for CFT.

Draft 1.1
9 sep 1998
All:

Fig. 1a and 1b:

All:
Font and format change.

Global change, figure numbers changed

Changes highlighted, covering: 

Letter 18feb1998, 5114/CP/LT/STJ-MAH, noted @@

VLT-VTE-98/0181, 30 jun 1998, noted: **

VLT-TRE-AMO-15100-0003, iss1, 30 jun 1998, noted: $$

VLT-VTE-98/0198, 14jul1998, noted: %%

VLT-MIN-ESO-15100-1625, noted: ##

Open issues: 

Station plug types.

Interface with Nasm wheel, Coude beam sw., FSS filter, x, y.

RS232 protocols.

Heidenhain 11uA or 1V system

Front space for ATS<>LCU connections

Volume of workstation

2
21 oct 1998
All:


Validated revisions as of draft 1.1;

Implemented changes as per att. 2 of VLT-MIN-ESO-15100-1669 (meeting 15-16 sep 1998);

Interface to Nasmyth Wheel and FSS equipment as in Issue. 1

Implemented changes as per VLT-AMO-98/0030 (15oct1998) and answer VLT-VTE-98/0298

3

prep. 1
20 feb 1999
All:
Restructured paragraphs layout.

Added all new interfaces as identified on SW Requirements Specification document VLT-SPE-ESO-15151-1712.

Changed accordingly to AMOS Preliminary Design Review documents.

3

prep. 2
10 mar 1999
All
Updated according to AMOS comments on prep.1

Restructured all paragraphs for better readability

Added Telescope Temperature Monitor to Azimuth LCU

Moved Nasmyth Wheel device from Azimuth to Altitude LCU

3
15 mar 1999

Moved FSS Filter Wheel, Translation Stage to Sensor LCU

Added FSS Field Diaphragm in Sensor LCU

Foresee Atmospheric Dispersion Corrector on Azimuth LCU

Added signal reference for brakes in Alt/Az drives

Added PHASE encoder signals input

3.1
30 mar 1999

Updated according to AMOS Fax VLT-AMO-99/0017

4.0
30 jul 1999

Updated according to :

AMOS Fax VLT-AMO-99/0037

ESO Fax VLT-VTE-98/0222, VLT-VTE-98/0279,VIF-99/0017 and VLT-VTE-99/0056

Minutes of meeting VLT-MIN-ESO-15100-1861 

Added future subsystems M10DF, M6DM, TADC

New Interlock System Motion Stop

4.1
10 Sep 1999

Updated according to :

Minutes of meeting VLT-MIN-ESO-15100-1910 

Fax VLT-AMO-99/0076

4.2
20 Sep 2000
Alt & Az









Include paragraph with co-ordinates conventions for Alt & Az axes
Phase amplifier: §3.2.2 & 3.3.2
take into account the ICD AMOS/PHASE  Issue 1

item 19 Convert ‘Drive ready ‘ into ‘Motor_Powered’
merge 50 &70 overtemp item 17&18 into No_Fault signal, the origin of the fault is retrieved via serial link.  
Tacho OK: High = Tacho OK; encoder is used for commutation
The torque is the monitored variable by default (tbc) p21&35
Item 4&5 Inhibit to ATS are accepted only when amplifier is in speed mode (i.e. maintenance or tests)



Interlock 
ref: TEC-TES-99/038













chapter Altitude 3.2.2 & Azimuth3.3.2;
add item 94  'BRAKE DISENGAGE' from LCU to ATS, active when contact closed.
Description: activates the brake Electro-pneumatic valve to disengage the brakes.

–chapter  LCU  Azimuth3.3.2;
item 86  change the naming to 'Transporter relocation' from ATS to LCU, active when contact open. (already defined in 3.5.5 item 11)
Description: action Az immobilized because AT not in station position  , connection to X15_nn 

add chapter3.5.5; item 12'Motion Stop Status' from ATS to LCU description, give the list of Motion Stop buttons currently engaged via the RS232 link. 

Add chapter Altitude 3.2.2; item 83 and chapter 3.3.2 Azimuth Item 86 change the naming to 'Transporter relocation' from ATS to LCU, active when contact open. (already defined in 3.5.5 item 11)
Description: action Alt immobilized because AT not in station position  , connection to X11_nn

Add chapter 3.3.2 Azimuth: item 88 'Enclosure  area' from ATS to LCU, active when contact open. Description: Azimuth AND Enclosure interlocked while maintenance is being performed inside the enclosure.

Chapter 3.3.2 Azimuth; item 83, 84 change naming ‘AT Motion Stop’ into 'Motion Stop' for consistency with item 80-81 of 3.2.2 Altitude. 

Chapter 4 Interlocks 
-add item number
- 'Enclosure area' is preferred to 'Altitude Maintenance' 

Chapter 3.1.4 Station Plug
item 13 Motion Stop,  1 contacts to Altitude, 1 contact to Azimuth, free of potential 
item 14 from: ATS to: LCU; Signal: Transporter relocation contact free of potential (the Carriage interlock; is AMOS internal

Overview schematic p 13 
-name: ‘VLTI Motion Stop’ instead of ‘AT Motion Stop’ 
-the VLTI Motion Stop goes to AMOS Motion Stop Chain 
-the carriage lock goes to Contractor carriage controller (Siemens PLC)

P20 item 64 contact closed if Altitude Locking pin DISENGAGED

P34 Azimuth direction switch item 62&63 add contact with opposite logicitem 62&63

§3.5.5 item 11 Relocation status two contacts instead of a single one; 
-Relocation to Altitude contact free of potential from ATS to LCU 
-Relocation to Azimuth contact free of potential from ATS to LCU

§3.5.5 add item 12 Motion Stop status; two contacts 
-Motion Stop to Altitude: contact free of potential from ATS to LCU 
-Motion Stop Azimuth: contact free of potential from ATS to LCU

the motion stop identification is done via the RS232 W42.

p39 chap3.3.5add after'…the guide pin.':
'when the limit switch is active azimuth is interlocked ; a manual override is provided in case M7 is in front of the GIS door and prevent maintenance access to the azimuth cable wrap.'



ICD
X naming of ICD terminals changed to Z



Field  Stabilisation System
Replace Strap TB Y32 by defined in AD03 connectors p45



FSS Translation stage
'….and a negative lmit switch(NLS).' Add 'One of them is used as reference switch.



M6
ICD requirements for the M6 drives (tip/tilt mirror and future AO)
Electrical power: 230V UPS 0.75mm2
Electrical signals: 1 x WAGO interface module  
Cooling connections
Fibre link to M6 and 12x2 electrical signals for control lines



Enclosure
The two Young anemometers are configured with analog output, reading of UVW is via the Siemens PLC. 



Station plug
The detailed Station Plug connections are given in  RD10
There 12 fibres 24 electrical contacts page 12 chap 3.1.4
§3.1.4 Item 13 is VLTI to ATS , Item 14 is ATS to ATS, Item 1&10 are Station Plug to ATS
Updated drawing



Azimuth Cable wrap
The azimuth axis shall be stopped whenever one of the proximity switches is activated. In case it stops the azimuth axis while M7 is blocking the access door; the cablewrap wrap can be manually dragged away. For that purpose the motorization drive is made reversible



Mech. Int.
§6.5.2 Dual feed LCU  the space required is defined in §6.1







20/09/00
M10
Power ON/OFF command LCU  to ATS OEM070 for reduced heat dissipation in the ROS. Digital


19/01/01
Z11 Terminal block 
§3.2.2 Include Z11 terminal block: Extension to Signal cabinmet via Az. Cable wrap  


19/01/01
Altitude Cabinet Temperature
§3.2.6 Include Temperature monitoring of  Altitude cabinet via serial port W12  


19/01/01
Annex01
Detail drawings LCU's Electrical and optical connections
§3.2.6 Include Temperature monitoring of  Altitude cabinet via serial port W12  


19/01/01
comments
Included comments  VLT-AMO-01/0006 
§3.5.5
- replaced Station plugged in by 400V OK and EG-Cmd-Off, 
-removed Transporter Ground Seal inflated
-status of pneumatic module replaced by Compressed air 
§3.5.2
-updated anemometer output
-item 104 to 107 replaced by the position readout of the four enclosure shells actuators
§3.2.6: Temperature signal is 4-20m; Full scale range is t.b.d.
§3.3.2:added CW positive & negative switches items 71to 76
§3.5.3: Included Dew point 
§4: Updated interlock table
§5: Electrical&Optical interface update
§6: Mechanical interfaces update

4.3
14 May 2001
comments
-Updated PHASE interface definition:
§3.2.2 Altitude drive and §3.3.2 Azimuth drive:added item 20 mode (torque/speed) updated item 30 Tacho signal, 
-§3.2.2 .1 Altitude Drive  Z11 is ESO Only 
-§3.2.6:   Telescope temperature sensors: 
item is Temperature Tube, item 7 is Temperature Center Piece

-Harmonised ICD naming of Maccon controlled motors between ICD, annex 01 and FISBA design report VLT-TRE-FIS-15100-043 issue 4.1. in §3.2.5 Nasmyth wheel, §3.3.4 Coude beam switching device, §3.4.6 FSS Filter Wheel, §3.4.7 FSS Aperture Fiel Diaphragm, §3.4.8, FSS Translation stage  
-§3.3.2 cable wrap limit switches description item 71 to 76  moved to §3.3.5. Azimuth cable wrap; Switches are electro-mechanical type.
-§3.4.6 FSS aperture corrected formatting; all is "ESO Only"
-§3.5.2 .1Enclosure: item 100 to 103 the open commands and close commands are replaced by open/close and start/stop commands 
-§3.5.2, §3.5.4, §3.5.6 Siemens PLC serial link is RS422 with protocol RK512/3964R
-§3.5.6 Service module: item 11 contact opens if battery NOT OK
-§4.1.1.1 Interlock Altitude Locking pin is received by Altitude drive only
-§6.1 added sentence for rear side access to Signal and Altitude cabinet 
 


TABLE OF CONTENTS

91
Scope

2
Documents and Acronyms
9
2.1
Applicable Documents
9
2.2
Reference Documents
9
2.3
Annex Documents
10
2.4
Acronyms
11
3
Interface Definition
12
3.1
Deployment of sub-systems on the LCUs
12
3.1.1
Introduction
12
3.1.2
Conventions
13
3.1.3
General requirements
13
3.1.4
Station Plug
15
3.2
Altitude LCU
18
3.2.1
LCU Configuration
18
3.2.2
Altitude drive
20
3.2.3
M2 drives
26
3.2.4
Pupil Beacon
29
3.2.5
Nasmyth wheel
30
3.2.6
Telescope Temperature Sensors
32
3.3
Azimuth LCU
34
3.3.1
LCU Configuration
34
3.3.2
Azimuth drive
36
3.3.3
M10 Drives
41
3.3.4
Coudé Beam Switching Device
42
3.3.5
Azimuth Cable Wrap
44
3.3.6
M10 Dual Feed Mirror
45
3.3.7
Transversal Atmospheric Dispersion Compensator
46
3.4
Sensor LCU
47
3.4.1
LCU Configuration
47
3.4.2
Field Acquisition System (FAS)
49
3.4.3
Field Stabilization System (FSS)
50
3.4.4
M6 drives
53
3.4.5
FSS Filter Wheel
55
3.4.6
FSS Aperture Field Diaphragm
57
3.4.7
FSS Translation Stage
59
3.4.8
M6 Deformable Mirror
62
3.5
Auxiliary LCU
63
3.5.1
LCU Configuration
63
3.5.2
Enclosure and Anemometers
Error! Bookmark not defined.
3.5.3
Air Conditioning
67
3.5.4
Relay Optics Structure (ROS) Shutter
Error! Bookmark not defined.
3.5.5
Transporter
Error! Bookmark not defined.
3.5.6
Service modules
Error! Bookmark not defined.
4
Interlocks
73
5
Electrical & Optical Interfaces
75
5.1
Station Plugs
75
5.2
Cables
75
5.3
PLC I/O
75
6
Mechanical Interfaces
77
6.1
Local Control Units
77
6.2
Detector and M6 drives
77
6.3
Pupil Beacon
77
6.4
Nasmyth focus
77
6.5
Future Extensions
77
6.5.1
Adaptive Optics
78
6.5.2
Dual Feed
78
6.5.3
Atmospheric Dispersion Corrector
78
6.6
Connection type & location
79
6.6.1
Terminal blocks
79
6.6.2
Serial interfaces
79
6.6.3
Heidenhain encoders
79
6.6.4
Optical fibre connection
79
6.6.5
ESO Terminal Blocks Yn
79
6.7
Cooling
79



Scope

This document describes the requirements and establishes the detailed specifications for the interface between the electro-mechanical hardware and the control system of the Auxiliary Telescope System. 

For the sake of completeness, this document includes also interfaces between the Control Software and the Control Hardware developed under ESO responsibility. 

In order to identify these interfaces internal to ESO, the associated section headers will be written in bold italic, the section content will be framed and marked with a special margin.

The document covers now the basic configuration (4 LCUs) which will be first implemented. Requirements associated to further extensions (e.g. future Adaptive Optics or Dual Feed) as defined in the ATS Technical Specifications (RD 01) and further clarified in correspondence have been included.  

<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Documents and Acronyms

Applicable Documents

The following documents, of the exact issue shown, form part of this ICD to the extend specified herein. In the event of conflict between the documents referenced herein and the content of this ICD, the content of this ICD shall be considered as a superseding requirement. 
  
  
  

PRIVATE
Reference 
Document Number 
Issue 
Date 
Title 

[AD 01]
VLT-SPE-ESO-10000-0015
4
 
Electronic design specifications

[AD 02]
VLT-SPE-ESO-10000-0003
1
05/02/1992
Electromagnetic compatibility and power quality specifications, Part 2

[AD 03]
VLT-ICD-ESO-11670-1288
3.0
21/12/1998
APD Tilt Loop Interfaces Control Document

[AD 04]
VLT-MAN-ESO-11670-1754
1.0
31/01/1999
STRAP Software Maintenance Manual

[AD 05]
PHASE AX_V
3.0
07/01/1999
Main user manual

[AD 06]
VLT-TRE-AMO-151000-028
3.0
25/06/1999
ATS Safety Assessment Report

[AD 07]
Heidenhain GmbH
-
02/1996
IK 320 VME-Bus Interface Board User Manual

[AD 08]
VLT-SPE-ESO-15151-1795
1.0
20/07/1999
Auxiliary Telescope Control Software 
System Design Description

Reference Documents

The following documents contain additional information and are referenced in the text: 
  
  
  

PRIVATE
Reference 
Document Number 
Issue 
Date 
Title 

[RD 01]
VLT-SPE-ESO-15400-0886
2.0
18/12/1996
VLTI Software Requirements Specifications

[RD 02]
VLT-SPE-ESO-15100-0338
2.1
19/05/1998
Technical Specifications for the ATS

[RD 03]
VLT-TRE-AMO-151800-082
2.0
28/05/1999
ATS AMOS Control System 
Final Design & Analysis Report

[RD 04]
VLT-TRE-FIS-151100-043
2.0
25/03/1999
ATS FISBA Package - Final Design Report

[RD 05]
VLT-MAN-ESO-17240-0672
1.5
19/11/1997
VLT Software - CCD Detector Control Software 
User Manual

[RD 06]
SIEI Peterlongo SpA
1.3
29/05/1995
S-Link 3 Protocol Specification

[RD 07]
VLT-MAN-ESO-11670-1870
1.0
in prep.
VLT Software - STRAP Control Software 
User Manual

[RD 08]
VLT-TRE-ESO-15154-1862
1.0
14/10/1999
Auxiliary Telescope System 
Control Hardware Design Specification 
Preliminary Design Report

[RD 09]
VLT-MAN-ESO-17210-0600
1.7
02/10/1998
VLT Software - Motor Control Software 
User Manual

[RD 10]
Sauter 
-
04/92
EY-2400 Protocol Specification

[RD 11]
VLT-SCH-ESO-15154-2202 
0
14/7/2000
AT Station Pit/Connection box internal wiring

[RD 12]
VLT-SPE-ESO-10000-0016 
2
07/10/1992
Basic Telescope definitions

[RD 13]
VLT MAN DJO 11700-0001
1
02/12/1991
TCCD System User's manual

[RD 14]
VLT-CRE-AMO-151240-086
3.0
10/04/2000
Station Plug Interface

[RD 15]
VLT-CRE-AMO-151123-057
2.0
08/05/2000
M2 Pupil Beacon Interface

[RD 16]
VLT-CRE-AMO-151281-124
1.0
05/06/2000
Dew point sensor inside enclosure

[RD 17]
VLT-CRE-AMO-151800-145
2.0
08/08/2000
Installation of ESO cables

[RD 18]
VLT-ICD-AMO-151800-137
1.0
08/03/2000
Interface control Document for the M2 support

Annex Documents

The following documents contain detailed interface drawings and are referenced in the text: 
  

PRIVATE
Reference
Document Number
Issue
Date
Title

[AN 01]
VLT-TRE-ESO-15154- 2400
1.1
22/01/2001
Auxiliary Telescope System LCU's 
Electrical & Optical connections 

Acronyms

This document employs several abbreviations and acronyms to refer concisely to an item, after it has been introduced. The following list is aimed to help the reader in recalling the extended meaning of each short expression: 
  

PRIVATE
ACE
Array Control Electronics

ADxx
Applicable Document #xx

ADC
Atmospheric Dispersion Corrector (Compensator)

AFD
Aperture Field Diaphragm

ALT
Altitude Axis

APD
Avalanche Photo-Diode

AT
Auxiliary Telescope unit

ATS
Auxiliary Telescope System

ATCS
Auxiliary Telescope Control System

CBSD
Coudé Beam Switching Device

DFM
Dual Feed Mirror

EMC
Electromagnetic Compatibility

ESO
European Southern Observatory

FAS
Field Acquisition System

FSS
Field Stabilization System

GIS
Ground Interface Structure

LAN
Local Area Network

LCM
Liquid Cooling Module

LCU
Local Control Unit

NDF
Neutral Density Filter

PE
Protective Earth

PLC
Programmable Logic Controller

ROS
Relay Optics Structure

RDxx
Reference Document #xx

STRAP
System for Tip-tilt Removal with APD

TADC
Transversal Atmospheric Dispersion Corrector (Compensator)

TCCD
Technical CCD System

TIM
Time Interface Module

WS
(Unix) Work Station

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 10:22:04 MET 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Interface Definition

Deployment of sub-systems on the LCUs

Introduction

The interfaces between the Control Software of the ATS and the Electro-Mechanical Hardware are implemented at LCU level. 

Typically one LCU is fully responsible for one or more sub-systems. 

<!-- ESO-ICD -->
PRIVATE
E 
S 
O 
When mapping the interfaces for the sub-systems described in the following sections with the requirements captured in the Use Cases, take into account that every sub-system corresponds to an Actor in the Use Cases.  

<!-- /ESO-ICD -->
The Auxiliary Telescopes control system is distributed over 1 Workstation and 4 separate Local Control Units; 
It is responsible for the control of 13 main and 5 auxiliary electro-mechanical functions, and 2 detectors. 

· Altitude LCU 

· Altitude drive 

· M2 drives 

· Pupil beacon 

· Nasmyth wheel 

· Temperature Sensors 

· Azimuth LCU 

· Azimuth drive 

· M10 (to be replaced by Dual Feed Mirror M10DFM) 

· Coudé Beam Switching Device 

· Azimuth cable wrap 

· M10 Dual Feed Mirror 

· Transversal Atmospheric Dispersion Compensator 

· Sensor LCU 

· Field Acquisition System (FAS) 

· Field Stabilization System (FSS) 

· M6 (to be replaced by Deformable Mirror M6DM) 

· FSS Filter Wheel 

· FSS Aperture Field Diaphragm 

· FSS Translation Stage 

· M6 Deformable Mirror 

· Auxiliary LCU 

· Enclosure 

· Air Conditioning 

· Relay Optics Structure (ROS) Shutter 

· Transporter 

· Service modules: 

· Auxiliary Power 

· Hydraulic and Pneumatic systems 

· Liquid Cooling Module 

Each individual LCU is contained in an electronic cabinet and has a mechanical and an electronic interface. The mechanical interface is the location and fixation on the transporter. The electrical interface are the commands and monitoring signal lines necessary for the operation of the entire system under control of the central control workstation. 

The contractor is responsible for the low level specification (encoder resolution, friction, etc.) which are necessary to meet the performance requirements as defined in the Technical Specifications. Even though the final position control of these elements is under ESO responsibility, the contractor shall verify performance requirements using his own verification control electronics, except for the azimuth and altitude loops. 

Conventions

All Terminal Blocks and connectors referenced in this document have been labelled as described below: 
  

PRIVATE
Label 
Connector Type 
Description 

Z 
Terminal Block 
Interface LCU to Contractor device 

W 
DB9 Connector 
Any RS232, RS422, or RS485 interface 

Y 
Terminal Block 
Interface LCU to ESO device 

The signals are described as seen from the LCU. 

The terminal blocks Yn and DB9 connectors Wn have been indexed following the product tree numbering implemented in [RD 08]: 
  

PRIVATE
LCU 
Index 
LCU 
Index 

Altitude 
1n 
Azimuth 
2n 

Sensor 
3n 
Auxiliary 
4n 

The LCU console connectors are indexed Wn0. 

General requirements

These interfaces shall be compliant with the Electronic Design Specifications [AD 01] and the Electro-Magnetic Compatibility Specifications [AD 02]. 

These interfaces shall be identical for every AT unit. 

The control racks shall be provided by the contractor (according to [AD 01]) and shall contain not only the power and drive electronics of the contractor but provide also the space for the ESO-LCUs. 

All 4 LCUs are connected to station plugs for the power, the time bus, the control LAN and the Guide / Acquisition LAN. 

All LCUs and other equipment supplied by ESO shall be powered by 230V via the battery system supplied by the contractor as shown in the following figure: 





Power distribution schematics

The electrical power of all this equipment is specified in the ATS technical specifications. 

LCU will be connected to UPS main power supply via circuit breakers and intermediate terminal blocks up to distribution bars generically named Z0 on the drawing.

  
  
  

Station Plug

The connection layout and pin assignment is given in Annex [AN1] AT Station Pit/Connection box VLT-SCH-ESO-15154-2202 
  

PRIVATE
Item 
From 
To 
Signal 
Description 


MAIN POWER: 






1 
Station plug 
ATS 
400VAC, 3 phases 
  


2 
  
  
Shield 
PE 


ELECTRICAL SIGNALS: 






10 
ATS 
Station plug 
+24 VDC 
  


11 
  
  
Shield 
PE 


STATION ID: 






12 
ATS 
LCU 
Station Id code 
8 Digital Input lines, pulled down to GND 
· upper 4-bit nipple r coding the rail=0x4r  
(in Hexadecimal: letters A to O)  

· lower nipple s coding the station=0x3s  
(in Hexadecimal: digits 0 to 9)  


AT MOTION STOP: 






13 
ATS 
LCU 
AT Motion Stop 
Contact free of potential: 
to Azimuth 
to Altitude 


OPTICAL FIBRES: 






20 
ATS 
LCU 
Control LAN 
1 fiber pair 


21 
ATS 
LCU 
Acquisition/Guide LAN 
1 fiber pair 


22 
ATS 
LCU 
Time Bus 
1 fiber pair 


23 
ATS 
LCU 
High Rate Data Links 
2 fiber pairs 


24 
ATS 
LCU 
Spare  
1 fiber pair 



  
  

Overview Schematic:





<!-- Standard footer -->

<!-- hhmts start -->Last modified: <!-- #BeginDate format:aEn2m -->Mon 14-May-2001 10:27<!-- #EndDate --><!-- hhmts end --> 




BLANK PAGE

<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Altitude LCU

LCU Configuration

<!-- ESO-ICD -->
  

PRIVATE
E 
S 
O 

O 
n 
l 
y 
Board 
Supplier 
Number 


CPU
Motorola MVME 2604 (PowerPC, 333 MHz, 16MB)
1 


Time Interface
ESO TIM
1 


Encoder interface
Heidenhain IK 320
1 


Amplifier Interface
ESO TBD
1 


Interlock Interface
ESO TBD
1 


Digital I/O Interface
Acromag AVME 9481
1 


Analog I/O Interface
VMIC VMIVME 3111
1 


Serial Interface
ESD ISER8
1 


Motion Controller
MACCON MAC4-INC
1 


Servo Amplifier
ESO VME4SA-01
1 


Hexapod controller
Delta Tau PMAC
1 

<!-- /ESO-ICD -->
This ICD specifies to use the couple ESO VME4SA servo amplifier (supplied by AMOS) and MACCON MAC4 Motion controller (supplied by ESO) for the implementation of the control of the Nasmyth wheel. 
The decision to assign the control of this wheel to this LCU was driven by the space requirement in the azimuth cable wrap. 
  
  

This LCU is responsible for: 
  

PRIVATE
Device 
Terminal Block 

Altitude drive
Z1,  Z11,  W13 

M2 drives
Z2,  W14 

Pupil Beacon
ESO Internal 

Nasmyth Wheel 
Nasmyth Beacon 
Z3 
ESO Internal 

Telescope Temperature Sensors
Z4, W12 

LCU Console
<!-- ESO-ICD -->  

PRIVATE
E 
S 
O 
The RS232 serial link is transiting via the Azimuth Cable Wrap to the Signal Cabinet on the DB9 Connector W10. see in Annex [AN01] VLT-DWG-ESO-15154-1769







Item 
From 
To 
Signal 
Description 
Pin 


LCU CONSOLE: 







1 
Altitude 
Cabinet 
Y10 
Signal 
Cabinet 
W10 
RS232 RX line
twisted pair
TBD 


2 


RS232 TX line




3 


RS232 GND line
 



4 


Shield
PE


<!-- /ESO-ICD -->
Schematic:

PRIVATE "TYPE=PICT;ALT=[Altitude LCU]"

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 10:47:31 MET 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Altitude drive

· Position control of the telescope altitude axis, including motor, amplifier, tachometer, encoder, brakes and limit switches. 

· The control is deployed on the Altitude LCU. 

· The device is connected to the Terminal Blocks Z1 & Z11, the later block being used to carry signals between Altitude and Signal cabinet via the Azimuth cable wrap 

· The amplifier is connected to the Connector W13. 

 The coordinate convention is defined in: Telescope basic definitions [RD 12] . The telescope tube points to zenith when the altitude angle a=90º, to horizontal direction when a=0º; 

The PID controller for the velocity of the Altitude drive will be included inside the LCU. It will provide an analogue signal as a reference with +-10Vdc maximum to the current loop, which belongs to the supply of the contractor and which may be included in his power amplifiers. 
The tachometer shall provide an analog signal (+-10Vdc) proportional to the actual speed of the drive system to the LCU as an analog symmetrical signal via a screened twisted pair cable. 
The connection between the encoder and the IK320 module in the LCU is defined in [AD 07]. 
The limit switch and brake status shall be connected to the ESO standard digital I/O board via the interlock system as defined in the Electronic Design Specification [AD 01]. The interlock system is supplied by ESO. The power 24VDC 0.6A for the PHASE drive is supplied by the LCU. 

<!-- ESO-ICD -->

  

<!-- /ESO-ICD -->
Altitude drive interfaces:

Proposed interface requirements according to the PHASE design using a fully digital amplifier, as in [AD 05]. 

The connection layout and pin assignment is given in [AN 01]Annex 1 VLT-DWG-ESO-15154-1769 
  
  
  

PRIVATE
Item 
From  
To  
Signal  
Description  

COMMUNICATION:  





 
 
 
PHASE Communication 
RS485 Port - Protocol S-Link 3 [RD 06]

AMPLIFIER:  





1 
LCU 
ATS 
Power On + 
Main power input relay
24VDC, 100mA supply to relay coil 
Digital output

2 
LCU 
ATS 
Power On - 
Common 

3 
LCU 
ATS 
Amplifier Enable 
High = Enable  
Contact closed to enable the amplifier output stage 
Toggling that signal resets externally triggered faults .e.i. PTC overtemperatures, Overspeed, encoder fail signal. Digital

4 
LCU  
ATS 

inhibit positive 
High = inhibit is asserted when positive vicinity switch is active. 
When the amplifier is in speed mode(manual only), the Positive velocity reference is clipped; the amplifier stop further motion in the corresponding direction.
When the amplifier is in torque mode (nominal configuration) the ESO velocity speed loop sends apropriate command to stop further motion in the corresponding direction.

5 
LCU 

ATS 
 
inhibit negative 
High = inhibit is asserted when negative vicinity switch is active.
When the amplifier is in speed mode(manual only), the Negative velocity reference is clipped; the amplifier stop further motion in the corresponding direction. 

When the amplifier is in torque mode (nominal configuration) the ESO velocity speed loop sends apropriate command to stop further motion in the corresponding direction

6 
ATS 
LCU 
Overspeed
High = no overspeed. The speed is sensed on the frequency of the Linear Hall Sensors, exceeding the limit will set the Overspeed alarm and keep it latched until the next enable cycle. 
Source Driver: 24V, 100mA max. Digital

7 
ATS 
LCU 
DriveOK (no-fault) 

High = DriveOK (no-fault). When negated (fault), the drive is disabled and latched until a reset is applied. Detailed fault information via the serial link. (Possible faults could be: tacho, hall sensor, over-under voltage, over current,, 50 &70 over temperature etc…)  
Source Driver: 24V, 100mA max. Digital

8 
ATS 
LCU 
400VAC Power OK 
High = the 400VAC Power is OK. 
Source Driver: 24V, 100mA max. Digital

9 
LCU 
ATS 
Amplifier drive power + 
+24VDC, 1.2A 

10 
LCU 
ATS 
Amplifier drive power - 
GND 

11 
 
 
unused 
PE 

14 
LCU 
ATS 
Torque command + 
Analog differential +/-10V, Zin = 10kW 
Analog output

15 
LCU 
ATS 
Torque command - 
Analog differential +/-10V, Zin = 10kW 
Analog output

16 
 
 
Shield
PE 

18 
LCU 
ATS 
Speed /Torque mode 
High = Speed mode.
The speed mode is used during manual operation for commissioning or maintenance, with a dedicated hand set. 
The torque mode is the nominal operation mode, the drive amplifier generates the tacho signal and uses the axis incremental encoder for improved ripple performances.

19 
ATS 
LCU 
Motor Powered 
High = 400V OK and Drive enabled 
Source Driver: 24V, 100mA max. 
Digital

TACHO:  





30 
ATS 
LCU 
Tacho signal + 
Analog symmetric +/-10V, 5mA  
Analog input

31 
ATS 
LCU 
Tacho signal - 
Analog input 

33 
LCU 
ATS 
Tacho Range 8x 
High = Sensitivity is increased by factor 8. Contact status: 24VDC, 6.6kW to GND 
Digital

34 
ATS 
LCU 
Tacho Range 
High Sensitivity is increased by factor 8.  
Source Driver: 24V, 100mA max. 
Digital,

35
ATS
LCU
Tacho OK
High = Tacho OK; the tacho signal is ouput to ESO LCU. 
When Encoder OK is asserted, the drive amplifier uses the sine & cosine signals from the axis encoder for commutation and improved ripple performances. If coherency between the two signals is lost the tacho OK is negated.
Source Driver: 24V, 100mA max. 
Digital, 

ENCODER:  1 head RON 905, 11mA 
Heidenhain IK320 output : 1 V peak-to-peak, sine-cosine A,B signals 





40 
LCU 
ATS 
A+ 
Output A+ 

41 
LCU 
ATS 
A- 
Output A- 

42 
 
 
Shield A 
 

43 
LCU 
ATS 
B+ 
Output B+ 

44 
LCU 
ATS 
B- 
Output B- 

45 
 
 
Shield B 
 

46 
 
 
Global Shield 
Chassis GND 

47 
LCU 
ATS 
Encoder OK 
High = Encoder OK and available,
The encoder is powered and provides the sincos incremental signals. The drive amplifier performs a coherency check between the sine and cosine signals. 
24VDC,Source Driver: 24V, 100mA max. 
Digital, 

SWITCHES:  





50 
ATS 
LCU 
Vicinity PLS + 
Contact opens if positive vicinity limit is crossed 
Digital 

51 
ATS 
LCU 
Vicinity PLS - 
Common 

52 
 
 
Shield 
PE 

53 
ATS 
LCU 
Vicinity NLS + 
Contact opens if negative vicinity limit is crossed 
Digital 

54 
ATS 
LCU 
Vicinity NLS - 
Common 

55 
 
 
Shield 
PE 

56 
ATS 
LCU 
Interlock PLS +  (Positive Limit+)
Contact opens if positive interlock limit is crossed 
Digital 

57 
ATS 
LCU 
Interlock PLS -  (Positive Limit-)
Common 

58 
 
 
Shield 
PE 

59 
ATS 
LCU 
Interlock NLS +  (Negative Limit+)
Contact opens if negative interlock limit is crossed 
Digital 

60 
ATS 
LCU 
Interlock NLS -  (Negative Limit-)
Common 

61 
 
 
Shield 
PE 

62 
ATS 
LCU 
Altitude Locking pinA + 
10 to 30 VDC 

63 
ATS 
LCU 
Altitude Locking pinA - 
Common 

64 
ATS 
LCU 
Altitude Locking pinA signal 
High = Altitude Locking pin is disengaged. 
Open collector PNP, 200 mA max. 
Digital

65 
ATS 
LCU 
Altitude Home switch + 
24VDC, 8W max. 

66 
ATS 
LCU 
Altitude Home switch - 
Common 

67 
 
 
Shield 
 

68 
ATS 
LCU 
Altitude Locking pinB + 
10 to 30 VDC 

69 
ATS 
LCU 
Altitude Locking pinB - 
Common 

70 
ATS 
LCU 
Altitude Locking pinB signal 
High = Altitude Locking pin is disengaged. 
Open collector PNP, 200 mA max. 
Digital

71 
ATS 
LCU 
EndStopA + 
On Nasmyth A side a Zenith or a Parking End Stop schock absorber is not in position Contact opens

72 
ATS 
LCU 
EndStopA - 
 Common

73 
ATS 
LCU 
EndStopB + 
On Nasmyth A side a Zenith or a Parking End Stop schock absorber is not in position Contact opens

74 
ATS 
LCU 
EndStopB - 
Common

80 
ATS 
LCU 
Motion Stop + 
24VDC, 8W max. 

81 
ATS 
LCU 
Motion Stop - 
Common 

82 
 
 
Shield 
 

83 
ATS 
LCU 
Transporter relocation+ 
High = Transporter in station position.  
Low = Altitude immobilized because AT not in station position.  
See IfTransporter_RelocationStatus 
1 Digital line, pulled down to GND

84 
 ATS
 LCU
Transporter relocation- 
 

85 
 
 
Shield 
 

BRAKES:  





90 
ATS 
LCU 
Status encoder side + 
8 NO switches in serie, 30 VDC, 10A max. 
Contact closed when brakes are disengaged 
Digital input

91 
ATS 
LCU 
Status encoder side - 
8 NO switches in serie, 30 VDC, 10A max. 
Contact closed when brakes are disengaged 
Digital input

92 
ATS 
LCU 
Status M4 side + 
8 NO switches in serie, 30 VDC, 10A max. 
Contact closed when brakes are disengaged 
Digital input

93 
ATS 
LCU 
Status M4 side - 
8 NO switches in serie, 30 VDC, 10A max. 
Contact closed when brakes are disengaged 
Digital input

94 
LCU 
ATS 
Brake command 
High = Enable  
Contact closes to disengage the brakes 

95 
LCU 
ATS 
Return 
Common 

MONITOR:  





100 
101
ATS 
LCU 
Signal monitoring 
Buffer amplifier output +/- 10V ,5mA differential for monitoring of internal amplifier variable selected via the serial communication link. 
Relevant signals for monitoring: 

· analog input   

· measured torque (default selection)   

· measured currents: phase 1, phase2, etc...   

· hall sensor signal   

· actual speed from the hall sensors   

· actual speed from the tacho   

· any signal specific to the amplifier/motor unit   

Analog input 

110 
LCU 
ATS 
Signal selection 
Saturation parameters: 

· maximum torque: the output torque cannot go beyond the set value  

· maximum speed: the output speed cannot go beyond the set value  

Controller parameters: 

· P,I,D servo loop gain adjustable on the fly  

· integrator anti windup  

· integrator limitation  


  



  

PRIVATE
Z11 TERMINAL BLOCK:AZIMUTH CABLEWRAP  






E 
S 
O

o 
n 
l 
y 

Item 
From 
To 
Signal 
Description 


1 
ATS 
LCU 
Motion Stop + 
24VDC, 8W max. 


2 
ATS 
LCU 
Motion Stop - 
Common 


3 
ATS 
LCU 
Relocation + 
High = Transporter in station position.  
Low = Altitude immobilized because AT not in station position. 


4 
ATS 
LCU 
Relocation - 



5-23 
SPARE 
SPARE 
 
 


24 
LCU 
LCU 
PHASE R/TX+ 
PHASE RS485 communication link 


25 
LCU 
LCU 
PHASE R/TX- 



26
LCU 
LCU 
PHASE return R/TX+ 



27
LCU 
LCU 
PHASE return R/TX- 



28
 
 
Ground 



29 
 
 
Ground 
PMAC RS232 communication link 


30 
LCU 
LCU 
PMAC TxD 



31 
LCU 
LCU 
PMAC RxD- 



32 
LCU 
LCU 
PMAC RTS 



33 
LCU 
LCU 
PMAC CTS 



34 
LCU 
LCU 
Ground 



35 
LCU 
LCU 
PMAC return TxD



36 
LCU 
LCU 
PMAC return RxD 



37 
LCU 
LCU 
PMAC return RTS 



38 
LCU 
LCU 
PMAC return CTS 



39 
LCU 
LCU 
CPU0 RXD 
RS232 communication link 


40 
LCU 
LCU 
CPU0 TXD 



41 
LCU 
LCU 
Ground 



42 
LCU 
LCU 
Ground 



43 
LCU 
LCU 
CPU1 TXD 
RS232 communication link 


44 
LCU 
LCU 
CPU1 RXD 



45-48 
 
 
SPARE 
 


  
  

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 10:59:26 MET 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
M2 drives

· Control of the focusing, centering and tilt mechanism (focus, x, y, a, b). 

· The control is deployed on the Altitude LCU. 

· The device is connected to the PMAC connectors J4, JS1, J8, JMACH1 & JMACH2. 

· The M2 controller is connected to the Connector W14. 

The M2 controller PMAC-VME is supplied by ESO. 

The connection layout and pin assignment is given in Annex 1 VLT-DWG-ESO-15154-2217

The control of M2 shall be performed by a PMAC VME board. The control of the PMAC board shall be done through RS422 interface on the ISER board. AMOS provides the source code to be implemented on the CPU in order to transform high-level commands into low-level commands (individual actuator motions). These commands are sent to the PMAC board via the RS422 interface. The high-level commands are listed in the table below. 

Important Note: The effective interface to the M2 drives is not on PMAC connectors J4, JS1, J8, JMACH1 & JMACH2. The contractor is responsible of the interface between the PMAC-VME output connector and the Terminal Block, while ESO is responsible of interfacing the PMAC-VME RS422 input port with a ESO standard serial ISER8 output port. see also the full description in RD 18 Interface control document for the M2 support. 

M2 drives interfaces:

PRIVATE
Item 
From 
To 
Signal 
Description 
Pin 

COMMUNICATION: 






1 
  
  
M2 Communication 
RS422 Port to PMAC Input port - Protocol details TBD 
W14 

GENERAL COMMANDS: 






10 
LCU 
ATS 
Power On/Off 
M2_POWER_ON_OFF 
Software Command to W14 

11 
LCU 
ATS 
Stop all actuators 
M2_STOP 


11.1 
LCU 
ATS 
Execute set command 
M2_GO 


12 
LCU 
ATS 
Constraints for each degree of freedom (focus, tip, tilt, x, y) 
M2_PARAMETERS min,max,velocity 


13 
LCU 
ATS 
Initialize all actuators 
(motion to Home position) 
M2_HOME 


14 
LCU 
ATS 
Request M2 status 
M2_STATUS_REQUEST
· M2_STATUS  

· M2_GLOBAL: 4-bit value indicating Power On/Off, home, moving and error  

· M2_FAULT_LOCATION: On error, indicates the faulty actuator (1 to 6)  

· M2_FAULT_DIAGNOSTIC: On error, indicates the error type  

· M2_TEMPERATURE_STATUS  

· SENSOR_TEMP: temperature measured by the sensor  


MOTION COMMANDS: 






20 
LCU 
ATS 
Set M2 absolute position 
M2_MOVE_ABS <axis <absolute value 
· M2_DES_FOCUS <z along Z axis in mm  

· M2_DES_TIP <a around X axis in radians  

· M2_DES_TILT <b around Y axis in radians  

· M2_DES_X <x along X axis in mm  

· M2_DES_Y <y along Y axis in mm  
Software Command to W14 

21 
LCU 
ATS 
Set M2 relative position 
M2_MOVE_REL <axis <relative value 
· M2_DES_FOCUS <z along Z axis in mm  

· M2_DES_TIP <a around X axis in radians  

· M2_DES_TILT <b around Y axis in radians  

· M2_DES_X <x along X axis in mm  

· M2_DES_Y <y along Y axis in mm  


POSITION COMMANDS: 






22 
LCU 
ATS 
Get M2 absolute position 
M2_MOVE_ABS <axis 
· M2_DES_FOCUS returns the position <z along Z axis in mm  

· M2_DES_TIP returns the position <a around X axis in radians  

· M2_DES_TILT returns the position <b around Y axis in radians  

· M2_DES_X returns the position <x along X axis in mm  

· M2_DES_Y returns the position <y along Y axis in mm  
Software Command to W14 

LOW LEVEL ACTUATOR CONTROL: 






30 
LCU 
ATS 
Set actuator absolute position 
M2_ACTUATOR_ABS <n=1..6 <absolute value 
· M2_DES_ACTUATOR_n_POSITION <p along its degree of freedom  
Software Command to W14 

31 
LCU 
ATS 
Set actuator relative position 
M2_ACTUATOR_REL <n=1..6 <relative value 
· M2_DES_ACTUATOR_n_POSITION <p along its degree of freedom  


32 
LCU 
ATS 
Get actuator absolute position 
M2_ACTUATOR_ABS <n=1..6 
· M2_DES_ACTUATOR_n_POSITION returns the position <p of actuator #n along its degree of freedom  



  
  
  <div align=right>
<!-- Standard footer --></div>

<!-- hhmts start -->Last modified: Mon Dec 18 10:20:37 MET 2000 <!-- hhmts end -->



Pupil Beacon
<!-- ESO-ICD -->
PRIVATE

E
S
O

O
n
l
y


· Control of the 2 pupil beacon light sources, infrared and visible fed into one optical fibre to the M2 beacon. 


· The control is deployed on the Altitude LCU. 

Pupil Beacon interfaces:






PRIVATE
Item
From
To
Signal
Description


PUPIL BEACON LIGHT SOURCES:






1
LCU
ATS
Switch Pupil beacon visible light source on/off
contact closed to switch on visible light source
Digital Output


2
ATS
LCU
Visible light source status
contact closed means visible light source switched on
Digital Input


3
LCU
ATS
Switch Pupil beacon infrared light source on/off
contact closed to switch on infrared light source
Digital Output


4
ATS
LCU
Infrared light source status
contact closed means infrared light source switched on
Digital Input

<!-- /ESO-ICD -->
<!-- Standard footer -->

<!-- hhmts start -->Last modified: Wed Sep 20 15:24:25 METDST 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Nasmyth wheel

· Control of the position of the wheel, including motor, tachometer, encoder and limit switches. This 4-element device consists of a Flat Retro-Reflecting mirror, a Free Hole and 2 positions for dedicated Alignment Tools (2 light beacons, half-masks). 
The motor and tacho are connected to the ESO standard VME4SA Servo Amplifier and the encoder to the MAC4 Motion Controller. The tacho generator is mounted on the motor. The velocity loop is closed in the amplifier. There are two switches: one of them is used as a reference switch. 

<!-- ESO-ICD -->
PRIVATE
E 
S 
O 
· Control of the 2 Image Beacon Light Sources in one of the dedicated Alignment Tools  
(they are fed into a single optical fibre). 

<!-- /ESO-ICD -->
· The control is deployed on the Altitude LCU. 

· The device is connected to the Terminal Block Z3. 

<!-- ESO-ICD -->
PRIVATE
E 
S 
O 

O 
n 
l 
y 
The motor control is a pure SW interface: the target position and velocity are passed to the Motion Controller via the VME bus. The motor control software is provided by ESO.  
The functionalities to be provided are:  

Set absolute position of Nasmyth wheel axis  

· Set to Retro-Reflecting mirror  

· Set to free hole  

· Set 1st Alignment Tool position (Beacon)  

· Set 2nd Alignment Tool position (Halfmasks)  

· Chop Nasmyth wheel with a given stroke and frequency  

<!-- /ESO-ICD -->
Nasmyth wheel interfaces:

Item 
From 
Signal 
Description 

MOTOR: Minimotor DC Brush 3557-024CS 
24VDC, 30W 





1
LCU
Motor M+
VME4SA Servo Amplifier Channel #1


2
LCU
Motor M-
VME4SA Servo Amplifier Channel #1


3
 
Shield MSH
 


TACHOGENERATOR: Minimotor 4.3 G60 
4.3mV / rpm 





4
ATS
Tacho T+
VME4SA Servo Amplifier Channel #1


5
ATS
Tacho T-
VME4SA Servo Amplifier Channel #1


6
 
Shield TSH
 


ENCODER: Minimotor Optical HP HEDL 5540-500 
5 VDC - 500 counts/turn - Resolution = 250 counts/degree = 14.4"/count 





7
LCU
Encoder 5V VCC
MAC4-INC Motion Controller Channel #1


8
LCU
Encoder GND
MAC4-INC Motion Controller Channel #1


10
ATS
Encoder A+
Differential line driver A+ 
MAC4-INC Motion Controller Channel #1


11
ATS
Encoder A-
Differential line driver A- 
MAC4-INC Motion Controller Channel #1


13
ATS
Encoder B+
Differential line driver B+ 
MAC4-INC Motion Controller Channel #1


14
ATS
Encoder B-
Differential line driver B- 
MAC4-INC Motion Controller Channel #1


16
ATS
Encoder I+
Differential line driver I+ 
MAC4-INC Motion Controller Channel #1


17
ATS
Encoder I-
Differential line driver I- 
MAC4-INC Motion Controller Channel #1


18
 
Shield
 


SWITCHES: Micromat KS35A11 
24V, 4A max. - Repeatability=0.03 degrees 





19
ATS
Reference & Negative Limit (NC)
MAC4-INC Motion Controller Channel #1


20
ATS
Reference & Negative Limit (COM)
MAC4-INC Motion Controller Channel #1


21
 
Shield 
 


22
ATS
Positive Limit (NC)
MAC4-INC Motion Controller Channel #1


23
ATS
Positive Limit Switch (COM)
MAC4-INC Motion Controller Channel #1


24
 
Shield 
 


25
 
PE 
 


26
 
PE 
 



  

Nasmyth Beacon light sources interfaces:

<!-- ESO-ICD -->
PRIVATE
E 
S 
O 

O 
n 
l 
y 
Item 
From 
To 
Signal 
Description 


IMAGE BEACON LIGHT SOURCES: 






100
LCU
ATS
Switch Image beacon light source #1 on/off
contact closed to switch on image beacon light source #1 
Digital Output


101
ATS
LCU
Image beacon light source #1 status
contact closed means light source #1 switched on 
Digital Input


102
LCU
ATS
Switch Image beacon light source #2 on/off
contact closed to switch on image beacon light source #2 
Digital Output


103
ATS
LCU
Image beacon light source #2 status
contact closed means light source #2 switched on 
Digital Input

<!-- /ESO-ICD -->
<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 10:20:44 MET 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Telescope Temperature Sensors

· Monitoring of the telescope temperature sensors 

· Monitoring of the Altitude cabinet temperature sensors 

· The control is deployed on the Altitude LCU. 

· The Telescope temperature sensors are connected to the Terminal Block Z4. 

· The Altitude cabinet Temperature sensors are connected to the connector W12 

Nine (9) temperature sensors PT100 are mounted on various parts of the telescope and are connected to signal conditioners providing a 4-20mA signal output. These signals are connected to a dedicated field interface (ESO supply) for signal formatting before entering the Analog Input ports. 
The sensors are monitoring the temperature on the following parts: 

· 2 on the telescope tube 

· 1 on the top ring 

· 2 on the fork 

· 1 on the mirror M7 

· 3 on the primary mirror M1 


The thermal controller PKC from EROELELECTRONIC keeps the Altitude cabinet internal temperature equal to the ambient air temperature. It acts on a motorized valve regulating the coolant flow inside the heat exchanger of the Altitude cabinet. One PT100 senses the ambient air temperature, the other one senses the Altitude cabinet internal air temperature. The LCU has got no control on the regulation but monitors the two temperature sensors connected to the thermal controller PKC. 

· 1 on the altitude cabinet 

· 1 on the ambient air 

Telescope Temperature Sensors Interfaces:

The connection layout and pin assignment is given in Annex 1 VLT-DWG-ESO-15154-1769 

PRIVATE
Item 
From 
To 
Signal 
Description 

TOP RING SENSOR: 





1
ATS
LCU
Temperature TopRing +
Current Loop 4-20mA (-10,+25C)

2
ATS
LCU
Temperature TopRing -
 

3
 
 
Shield 
 

TUBE SENSORS: 





4
ATS
LCU
Temperature Tube +
Current Loop 4-20mA (-10,+25C)

5
ATS
LCU
Temperature Tube -
 

6
 
 
Shield 
 

7
ATS
LCU
Temperature Center Piece +
Current Loop 4-20mA (-10,+25C)

8
ATS
LCU
Temperature Center Piece -
 

9
 
 
Shield
 

FORK SENSORS: 





10
ATS
LCU
Temperature Fork A +
Current Loop 4-20mA (-10,+25C)

11
ATS
LCU
Temperature Fork A -
 

12
 
 
Shield
 

13
ATS
LCU
Temperature Fork B +
Current Loop 4-20mA (-10,+25C)

14
ATS
LCU
Temperature Fork B -
 

15
 
 
Shield
 

M7 SENSOR: 





16
ATS
LCU
Temperature M7 +
Current Loop 4-20mA (-10,+25C)

17
ATS
LCU
Temperature M7 -
 

18
 
 
Shield
 

M1 SENSORS: 





19
ATS
LCU
Temperature M1 A +
Current Loop 4-20mA (-10,+25C)

20
ATS
LCU
Temperature M1 A -
 

21
 
 
Shield
 

22
ATS
LCU
Temperature M1 B +
Current Loop 4-20mA (-10,+25C)

23
ATS
LCU
Temperature M1 B -
 

24
 
 
Shield
 

25
ATS
LCU
Temperature M1 C +
Current Loop 4-20mA (-10,+25C)

26
ATS
LCU
Temperature M1 C -
 

27
 
 
Shield
 

Power Supply: 





28
LCU
ATS
Power Supply +
24VDC / 600mA for the 9 Temperature Sensors PT100 and Signal conditioners

29
LCU
ATS
Power Supply -
 

30
 
 
Shield 
 

Altitude cabinet Thermal controller:

PRIVATE
Item 
From 
To 
Signal 
Description 

COMMUNICATION: 





1
ATS
LCU
Thermal controller Communication
RS485 Port to PKC EROELECTRONIC - Protocol details TBD 

GENERAL COMMANDS: 





2
ATS
LCU
Read temperatures
READ_THERMAL_CONTROLLER

 

<!-- Standard footer -->

<!-- hhmts start -->Last modified: <!-- #BeginDate format:aEn2m -->Wed 28-Mar-2001 14:47<!-- #EndDate --><!-- hhmts end --> 




Azimuth LCU

LCU Configuration

<!-- ESO-ICD -->
PRIVATE
E
S
O

O
n
l
y
Board
Supplier
Number


CPU
Motorola MVME 2604 (PowerPC, 333 MHz, 16MB)
1


Time Interface
ESO TIM
1


Encoder interface
Heidenhain IK 320
2


Amplifier Interface
ESO TBD
1


Interlock Interface
ESO TBD
1


Digital I/O Interface
Acromag AVME 9481
1


Analog I/O Interface
VMIC VMIVME 3111
1


Serial Interface
ESD ISER8
1


Motion Controller
MACCON MAC4-INC
1


Servo Amplifier
ESO VME4SA-01
1

<!-- /ESO-ICD -->
This ICD specifies to use the couple ESO VME4SA servo amplifier (supplied by AMOS) and MACCON MAC4 Motion controller (supplied by ESO) for the implementation of the control of the Coudé Beam Switching device and of the Transversal Atmospheric Dispersion Compensator (foreseen at a later stage as an ESO supply). 

 

This LCU is responsible for:

PRIVATE
Device 
Terminal Block 

Azimuth drive 
Z5, W21 

M10 (to be replaced by Dual Feed Mirror) 
Z18 - W22 

Coudé Beam Switching Device 
Z10 

Azimuth cable wrap 
Z5, W21 

<!-- ESO-ICD -->
PRIVATE
E
S
O
Future Device
Terminal Block


M10 Dual Feed Mirror
Dual Feed LCU 


Transversal Atmospheric Dispersion Compensator
Y20 

<!-- /ESO-ICD -->
Schematic:

PRIVATE "TYPE=PICT;ALT=[Azimuth LCU]"

 

<!-- Standard footer -->

<!-- hhmts start -->Last modified: <!-- #BeginDate format:En1 -->23-mar-01<!-- #EndDate --> <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Azimuth drive

· Position control of the telescope azimuth axis, including motor, amplifier, tachometer, encoder, brakes and limit switches. 

· The control is deployed on the Azimuth LCU. 

· The device is connected to the Terminal Blocks Z5. 

· The amplifier is connected to the Connector W21. 

  The coordinate convention is defined in: Basic Telescope definitions [RD 12] . The telescope azimuth angle is A= 0º; when the tube set horizontal is directed to the south-point. The amount of rotation is measured eastwards of the south-point. 

The PID controller for the velocity of the Azimuth drive will be included inside the LCU. It will provide an analogue signal as a reference with +-10Vdc maximum to the current loop, which belongs to the supply of the contractor and which may be included in his power amplifiers. 
The tachometer shall provide an analog signal (+-10Vdc) proportional to the actual speed of the drive system to the LCU as an analog symmetrical signal via a screened twisted pair cable. 
The connection between the encoder and the IK320 module in the LCU is defined in [AD 07]. 
The limit switch and brake status shall be connected to the ESO standard digital I/O board via the interlock system as defined in the Electronic Design Specification [AD 01]. The interlock system is supplied by ESO. The power 24VDC 1.2A for the PHASE drive is supplied by the LCU. 

<!-- /ESO-ICD -->
Azimuth drive interfaces:

Proposed interface requirements according to the PHASE design using a fully digital amplifier, as in [AD 05]: 

The connection layout and pin assignment is given in Annex 1 VLT-DWG-ESO-15154-1770 
  

PRIVATE
Item
From
To
Signal
Description

COMMUNICATION:





 
 
 
PHASE Communication
RS485 Port - Protocol S-Link 3 [RD 06]

AMPLIFIER: 





1
LCU
ATS
Power On + 
Main power input relay
24VDC, 100mA supply to relay coil 
Digital 

2
LCU
ATS
Power On -
Common

3 
LCU 
ATS 
Amplifier Enable 
High = Enable  
Contact closed to enable the amplifier output stage 
Toggling that signal resets externally triggered faults .e.i. PTC overtemperatures, Overspeed, encoder fail signal. Digital

4 
LCU  
ATS 

inhibit positive 
High = inhibit is asserted when positive vicinity switch is active. 
When the amplifier is in speed mode(manual only), the Positive velocity reference is clipped; the amplifier stop further motion in the corresponding direction.
When the amplifier is in torque mode (nominal configuration) the ESO velocity speed loop sends apropriate command to stop further motion in the corresponding direction.

5 
LCU 

ATS 
 
inhibit negative 
High = inhibit is asserted when negative vicinity switch is active.
When the amplifier is in speed mode(manual only), the Negative velocity reference is clipped; the amplifier stop further motion in the corresponding direction. 

When the amplifier is in torque mode (nominal configuration) the ESO velocity speed loop sends apropriate command to stop further motion in the corresponding direction

6
ATS
LCU
Overspeed
High = no overspeed. The speed is sensed on the frequency of the Linear Hall Sensors, exceeding the limit will set the Overspeed alarm and keep it latched until the next enable cycle. 
Source Driver: 24V, 100mA max. Digital

7
ATS
LCU
DriveOK (no-fault) 

High = DriveOK (no-fault). When negated (fault), the drive is disabled and latched until a reset is applied. Detailed fault information via the serial link. (Possible faults could be: tacho, hall sensor, over-under voltage, over current,, 50 &70 over temperature etc…)  
Source Driver: 24V, 100mA max. Digital

8
ATS
LCU
400VAC Power OK
High = the 400VAC Power is OK. 
Source Driver: 24V, 100mA max. Digital 

9
LCU
ATS
Amplifier drive power +
+24VDC, 1.2A

10
LCU
ATS
Amplifier drive power -
GND

14
LCU
ATS
Torque command +
Analog differential +/-10V, Zin = 10kW 
Analog output

15
LCU
ATS
Torque command -
Analog differential +/-10V, Zin = 10kW 
Analog output

16
 
 
Shield 
PE

18
LCU
ATS
Speed / Torque mode
High = Speed mode.
The speed mode is used during manual operation for commissioning or maintenance, with a dedicated hand set. 
The torque mode is the nominal operation mode, the drive amplifier generates the tacho signal and uses the axis incremental encoder for improved ripple performances.Source Driver: 24V, 100mA max. 
Digital

19
ATS
LCU
Motor Powered
High = 400V OK and Drive enabled 
Source Driver: 24V, 100mA max. 
Digital,

TACHO: 





30
ATS
LCU
Tacho signal +
Analog symmetric +/-10V, 5mA  
Analog input

31
ATS
LCU
Tacho signal -
Analog input

32
 
 
Shield 
PE

33
LCU
ATS
Tacho Range 8x
High = Sensitivity is increased by factor 8. Contact status: 24VDC, 6.6kW to GND 
Digital

34
ATS
LCU
Tacho Range 
High = Sensitivity is increased by factor 8 
Source Driver: 24V, 100mA max. 
Digital,

35
ATS
LCU
Tacho OK
High = Tacho OK; the tacho signal is ouput to ESO LCU. 
When Encoder OK is asserted, the drive amplifier uses the sine & cosine signals from the axis encoder are for commutation and improved ripple performances. If coherency between the two signals is lost the tacho OK is negated.
Source Driver: 24V, 100mA max. 
Digital, 

ENCODER: 4 heads ERA 780 T 
Heidenhain IK320 buffered output : 1 V peak-to-peak, sine-cosine A,B signals 





40
LCU
ATS
A+
Output A+

41
LCU
ATS
A-
Output A-

42
 
 
Shield 
 

43
LCU
ATS
B+
Output B+

44
LCU
ATS
B-
Output B-

45
 
 
Shield 
 

46
 
 
Global Shield 
Chassis GND

47 
LCU 
ATS 
Encoder OK 
High = Encoder OK and available,
The encoder is powered and provides the sincos incremental signals. The amplifier performs a coherency check between the sine and cosine signals. 
24VDC,Source Driver: 24V, 100mA max. 
Digital, 

SWITCHES: 





50
ATS
LCU
Vicinity PLS +
Contact opens if positive vicinity limit is crossed 


51
ATS
LCU
Vicinity PLS -
Common

52
 
 
Shield 
PE

53
ATS
LCU
Vicinity NLS +
Contact opens if negative vicinity limit is crossed 


54
ATS
LCU
Vicinity NLS -
Common

55
 
 
Shield 
PE

56
ATS
LCU
Interlock PLS + (Positive Limit +)
Contact opens if positive interlock limit is crossed 

57
ATS
LCU
Interlock PLS - (Positive Limit -)
Common

58
 
 
Shield 
PE

59
ATS
LCU
Interlock NLS + (Negative Limit +)
Contact opens if negative interlock limit is crossed 


60
ATS
LCU
Interlock NLS - (Negative Limit -)
Common

61
 
 
Shield 
PE

62
ATS
LCU
Direction switch 1+
24VDC, 10W max. 
Contact  

· Closed if in range -180 - +45 deg   

· Open if in range +90 - +360 deg  

Digital 

63
ATS
LCU
Direction switch 1-
Common

64
ATS
LCU
Direction switch 2+
24VDC, 10W max. 
Contact  

· Open if in range -180 - +45 deg   

· Closed if in range +90 - +360 deg  

Digital 

65
ATS
LCU
Direction switch 2-
Common

80
ATS
LCU
Interior door switch +
220VDC

81
ATS
LCU
Interior door switch -
220 VDC

82
 
 
Shield 
 

83
ATS
LCU
Motion Stop +
Contact open Motion Stop is active. 
24VDC, 8W max.

84
ATS
LCU
Motion Stop -
Common

85
ATS
LCU
Station Id code
8 Digital lines, pulled down to GND

86
ATS
LCU
Transporter relocation
High = Transporter in station position.  
Low = Azimuth immobilized because AT not in station position.  
See Transporter_RelocationStatus 
1 Digital line, pulled down to GND

87
 
 
Shield 
 

88
LCU
ATS
Enclosure Working Area +
24V field

89
ATS
LCU
Enclosure Working Area -
Low = maintenance is being performed inside the enclosure; Azimuth AND Enclosure interlocked  
1 Digital line, pulled down to GND

BRAKES: 





90
ATS
LCU
Status -U side +
12 NO switches in serie, 30 VDC, 10A max. 
Contact closed when brakes are disengaged 
Digital 

91
ATS
LCU
Status -U side -
12 NO switches in serie, 30 VDC, 10A max. 
Contact closed when brakes are disengaged 
Digital 

92
ATS
LCU
Status +U side +
12 NO switches in serie, 30 VDC, 10A max. 
Contact closed when brakes are disengaged 
Digital 

93
ATS
LCU
Status +U side -
12 NO switches in serie, 30 VDC, 10A max. 
Contact closed when brakes are disengaged 
Digital 

94
LCU
ATS
Brake command+
High = Enable  
Contact closes to disengage the brakes

95
LCU
ATS
return
Common

MONITOR: 





100 
101
ATS
LCU
Signal monitoring
Buffer amplifier output +/- 10V ,5mA differential for monitoring of internal amplifier variable selected via the serial communication link. 
Relevant signals for monitoring: 

· analog input  

· measured torque (default selection)  

· measured currents: phase 1, phase2, etc...  

· hall sensor signal  

· actual speed from the hall sensors  

· actual speed from the tacho  

· any signal specific to the amplifier/motor unit  

Analog input

110
LCU
ATS
Signal selection
Saturation parameters: 

· maximum torque: the output torque cannot go beyond the set value  

· maximum speed: the output speed cannot go beyond the set value  

Controller parameters: 

· P,I,D servo loop gain adjustable on the fly  

· integrator anti windup  

· integrator limitation   

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 11:24:13 MET 2000 <!-- hhmts end -->



M10 Drives

· Control of the two low frequency M10 tilt motions, required for output pupil position adjustment. Control of the Dual Feed mirrors (when implemented) for output pupil position adjustment as well as reference object selection and tracking. 

· The control is deployed on the Azimuth LCU. 

· The device is connected to a RS232 port via connector W22.
The two axes are driven by one OEM070 controller each. The two controllers are daisy-chained on the communication link. 
The two controllers are located in the ROS and switched OFF when not in use for reduced heat dissipation 

 

The protocol is described in OEM070 User Guide (see Annex 1, p. 23ff. of [RD 03]). The contractor is responsible for the bootstrap procedure, to be burned in EPROM. The a angle is managed by the controller #1, while the b angle from the controller #2. 

The conversion between encoder counts (after quadrature interpolation) and radians (or degrees) is given by a formula, given by the contractor. 

M10 interfaces:

The connection layout and pin assignment of W22 and Z18 is given in Annex 1 VLT-DWG-ESO-15154-1770

PRIVATE
Item 
From 
To 
Signal 
Description 
Pin 

COMMUNICATION: 






1 
  
  
M10 Communication 
RS232 Port - Protocol ASCII
9600 baud, 8 data bits, 1 stop bit, no parity
Encoder repeatability 0.0065 degrees, resolution 244 counts/degree 
W22 

TILTING: 






10 
LCU 
ATS 
Set M10 absolute (a,b) tilt angles 
Sets absolute M10 tilt angles (a,b) in encoder counts (two separate commands, one per axis) 
Software
Command
to W22 

11 
LCU 
ATS 
Set M10 relative (a,b) tilt angles 
Sets relative M10 tilt angles (a,b) in encoder counts (two separate commands, one per axis) 


12 
ATS 
LCU
Get M10 absolute (a,b) tilt angles 
Gets absolute M10 tilt angles (a,b) in encoder counts (two separate commands, one per axis) 


<div align="center">CONTROLLER: </div>






20
LCU
ATS 
Power On +
24VDC, 500mA supply to relay coil
 

21 
LCU 
ATS 
Power On - 
return 



<!-- hhmts start -->Last modified: Monday, 03/05/2001<!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Coudé Beam Switching Device

· Control of the selection of the different elements that can be inserted in the beam in front of the Coudé sensor. It consists of 5 elements, namely a Mirror reflecting the beam to FAS only, a Beam Splitter sending the beam to both FAS and FSS, a Free Hole to FSS only, a Light Stop that prevents the beam to reach neither FAS nor FSS and a FSS pupil viewer which enables to view the image of the pupil on the FSS APD using the FAS. 
The motor and tacho are connected to the ESO standard VME4SA Servo Amplifier and the encoder to the MAC4 Motion Controller. The tacho generator is mounted on the motor. The velocity loop is closed in the amplifier. There are two switches; one of them is used as a reference switch. 

· The control is deployed on the Azimuth LCU. 

· The device is connected to the Terminal Block Z10. 

<!-- ESO-ICD -->
PRIVATE
E 
S 
O 

O 
n 
l 
y 
The motor control is a pure SW interface: the target position and velocity are passed to the Motion Controller via the VME bus. The motor control software is provided by ESO.  
The functionalities to be provided are:  

Set absolute position of Coudé Beam Switching Device axis:  

· Set Light Stop  

· Set FAS only  

· Set FAS+FSS Beam Splitter  

· Set FSS only (Free Hole)  

· Set FSS pupil viewer  

<!-- /ESO-ICD -->
Coudé Beam Switching Device interfaces:

The connection layout and pin assignment of Z10 is given in Annex 1 VLT-DWG-ESO-15154-1770

PRIVATE
Item 
From 
To 
Signal 
Description 

MOTOR: Minimotor DC Brush 2233-024S 
24VDC, 2.53W 





1
LCU
ATS
Motor M+
VME4SA Servo Amplifier Channel #1

2
LCU
ATS
Motor M-
VME4SA Servo Amplifier Channel #1

3
 
 
Shield MSH
 

TACHOGENERATOR: Minimotor 1.5 G 
1.5mV / rpm 





4
ATS
LCU
Tacho T+
VME4SA Servo Amplifier Channel #1

5
ATS
LCU
Tacho T-
VME4SA Servo Amplifier Channel #1

6
 
 
Shield TSH
 

ENCODER: Minimotor Optical HP HEDL 5540-500 
5 VDC - 500 counts/turn - Resolution = 15000 counts/mm = 66.7mm/count 





7
LCU
ATS
Encoder 5V VCC
MAC4-INC Motion Controller Channel #1

8
LCU
ATS
Encoder GND
MAC4-INC Motion Controller Channel #1

7
ATS
LCU
Encoder A+
Differential line driver A+ 
MAC4-INC Motion Controller Channel #1

8
ATS
LCU
Encoder A-
Differential line driver A- 
MAC4-INC Motion Controller Channel #1

13
ATS
LCU
Encoder B+
Differential line driver B+ 
MAC4-INC Motion Controller Channel #1

14
ATS
LCU
Encoder B-
Differential line driver B- 
MAC4-INC Motion Controller Channel #1

16
ATS
LCU
Encoder I+
Differential line driver I+ 
MAC4-INC Motion Controller Channel #1

17
ATS
LCU
Encoder I-
Differential line driver I- 
MAC4-INC Motion Controller Channel #1

18
 
 
Shield 
 

SWITCHES: Micromat KS35A11 
24V, 4A max. - Repeatability=0.03mm 





19
ATS
LCU
Reference & Negative Limit(NC)
MAC4-INC Motion Controller Channel #1

20
ATS
LCU
Reference & Negative Limti (COM)
MAC4-INC Motion Controller Channel #1

21
 
 
Shield
 

22
ATS
LCU
Positive Limit (NC)
MAC4-INC Motion Controller Channel #1

23
ATS
LCU
Positive Limit (COM)
MAC4-INC Motion Controller Channel #1

24
 
 
Shield
 

25
 
 
Global Shield
 

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 11:25:32 MET 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Azimuth Cable Wrap

· The cable wrap control is an active motorized system and acts as a slave of the azimuth drive. No control by ESO is required but status information shall be provided as listed in the table below. 
The cable wrap guide has 2 electro-mechanical switches located at each side of the guide pin. 
The azimuth axis shall be stopped whenever one of the switches is activated. In case it stops the azimuth axis while M7 is blocking the GIS access door; the cablewrap wrap can be manually dragged away. For that purpose the motorization drive is made reversible 

· The control is deployed on the Azimuth LCU. 

· The device is connected to the Terminal Block Z5 and J6 

· The amplifier is connected to the Connector W21. 

Azimuth Cable Wrap interfaces:

The connection layout and pin assignment of Z5 and J6 is given in Annex [AN1] VLT-DWG-ESO-15154-1770

PRIVATE
Item 
From 
To 
Signal 
Description 

COMMUNICATION: 





 
 
 
PHASE Communication
RS485 Port - Protocol S-Link 3 [RD 06]

CABLE WRAP: 





1
ATS
LCU
No Fault
Contact closed if cable wrap OK (see interlock section) 
Source driver 24V, 100mA max. 
Digital Input

2
ATS
LCU
Torque signal + [Nm/V]
Analog signal -10V/+10V, 5mA (differential) 
Analog input

3
ATS
LCU
Torque signal -
Analog input

4
 
 
Shield
 

5
LCU
ATS
Amplifier Enable
Contact status 6.6kW to GND 
Digital Output

SWITCHES: 





8
ATS
LCU
Cable Wrap Positive Limit+
Contact opens if Cablewrap do not follow azimuth axis in positive direction

9
ATS
LCU
Cable Wrap Positive Limit-
contact return

10
ATS
LCU
Cable Wrap Negative Limit+
Contact opens if Cablewrap do not follow azimuth axis in negative direction

11
ATS
LCU
Cable Wrap Negative Limit-
return

<div align=right><!-- Standard footer --></div>

<!-- hhmts start -->Last modified: <!-- #BeginDate format:aEn2m -->Mon 07-May-2001 8:41<!-- #EndDate --><!-- hhmts end --> 




M10 Dual Feed Mirror
<!-- ESO-ICD -->
PRIVATE

E
S
O

O
n
l
y

 

This function will not be implemented in the initial phase of the project, but at a later stage. However interface requirements related to reserved space, cable routing, plug design, etc... apply. 

 

· Control of the M10 Dual Feed Mirror. 


· The control is deployed on the Dual Feed LCU. 

· This sub-system belongs to the ESO delivery. 

<!-- /ESO-ICD -->
<!-- Standard footer -->

<!-- hhmts start -->Last modified: Wed Aug 4 08:04:49 METDST 1999 <!-- hhmts end -->



Transversal Atmospheric Dispersion Compensator
 

<!-- ESO-ICD -->
PRIVATE

E
S
O

O
n
l
y

 

This function will not be implemented in the initial phase of the project, but at a later stage. However interface requirements related to reserved space, cable routing, plug design, etc... apply. 

 

· Control of the Transversal Atmospheric Dispersion Compensator. 
The motor and tacho are connected to the ESO standard VME4SA Servo Amplifier and the encoder to MAC4 Motion Controller. The tacho generator is mounted on the motor. The velocity loop is closed in the amplifier. There are two switches; one of them is used as a reference switch. 


· The control is deployed on the Azimuth LCU. 

· This sub-system belongs to the ESO delivery. 

· The device is connected to the Terminal Block Y20. 

Transversal Atmospheric Dispersion Compensator drive interfaces:

The pin assignment on Terminal Block Y20 will be identical to the ones used on the Coude Beam Splitting device. 

<!-- /ESO-ICD -->
The connection layout and pin assignment of Y20 is given in Annex 1 VLT-DWG-ESO-15154-1770

<!-- Standard footer -->

<!-- hhmts start -->Last modified:<!-- #BeginDate format:aEn2m -->Wed 13-Dec-2000 16:53<!-- #EndDate --><!-- hhmts end --> 




Sensor LCU

LCU Configuration

<!-- ESO-ICD -->
PRIVATE
E
S
O

O
n
l
y 
Board 
Supplier 
Number 


CPU 
Motorola MVME 2604 (PowerPC, 333 MHz, 32/64MB) 
1 


Time Interface 
ESO TIM 
1 


Digital I/O Interface 
Acromag AVME 9481 
1 


STRAP VME control board 
Microgate 
1 


TCCD boards 
LIB (INMOS BO16) + LAB 
1 


28V power supply for TCCD 
Kniel CP28.2 
1 


Motion Controller 
MACCON MAC4-INC 
1 


Servo Amplifier 
ESO VME4SA-01 
1 


+5V +/-15V Supply 
KNIEL CPM 101 
1 

<!-- /ESO-ICD -->
This ICD specifies to use the couple ESO VME4SA servo amplifier and MACCON MAC4-INC motion controller (both supplied by ESO) for the implementation of the control of the FSS Filter Wheel, FSS Field Diaphragm and the FSS Translation Stage X- & Y-axes. 

 

This LCU is responsible for:

PRIVATE
Device
Terminal Block

FSS Filter Wheel
Z12

FSS Translation Stage
Z16 & Z17

<!-- ESO-ICD -->
PRIVATE
E
S
O
The field acquisition system, the field stabilization system and the M6 drives and the Field Diaphragm as shown on the Figure below belong to the ESO delivery. There is for those parts only a mechanical interface existing (see 6.3).

<!-- /ESO-ICD -->
<!-- ESO-ICD -->
PRIVATE
E
S
O

O
n
l
y 
Device 
Connection 


Field Acquisition System (FAS) 
Y31 


Field Stabilization System (FSS) 
See [AD 03] 


M6 (to be replaced by Deformable Mirror) 
Y33 


FSS Field Diaphragm 
Y34 


Flow Meter 
Y35 


Power Supplies +5VDC +/-15VDC & 28VDC 
Y31 


Future Device 
  


M6 Deformable Mirror 
Adaptive Optics LCU 

<!-- /ESO-ICD -->
Cooling Monitor:

For safety and maintenance purpose, a flow-meter is installed in the liquid cooling circuit. This device can detect rapidly any leak or obstruction in the cooling pipes. 

Schematic:

PRIVATE "TYPE=PICT;ALT=[Sensor LCU]"

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 11:33:18 MET 2000 <!-- hhmts end -->



Field Acquisition System (FAS)
<!-- ESO-ICD -->
PRIVATE
E
S
O

O
n
l
y
· Control of the Field Acquisition System components (CCD head, pre-amplifier, temperature sensor, liquid cooling). This system is used for Auto Guiding. 


· The control is deployed on the Sensor LCU. 

· This sub-system belongs to the ESO delivery. 

· The device is connected to the Terminal Block Y31. 

Field Acquisition System interfaces:

The reference document for the FAS interface is [RD 05]. 

The connection layout and pin assignment of Y31 is given in Annex 1 VLT-DWG-ESO-15154-1787 






PRIVATE
Item
From
To
Signal
Description


TCCD POWER SUPPLY:






1
LCU
ATS
Power Line +
28VDC, 2A


2
LCU
ATS
Power Line -
 


3
 
 
Shield 
 


TCCD OPTICAL FIBRE:






13
LCU
ATS
Optical Fibre #n (n=1 to 8)
See [RD 05] and Note 1


 

Note 1: From LAB Board Front panel to ROS Plug 





<!-- /ESO-ICD -->
 

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 12:27:14 MET 2000 <!-- hhmts end -->



Field Stabilization System (FSS)

<!-- ESO-ICD -->
PRIVATE
E
S
O

O
n
l
y
· Control of the Field Stabilization System components (STRAP system, communication link to M6, liquid cooling). This system is used for Field Stabilization. 

· The control is deployed on the Sensor LCU. 

· This sub-system belongs to the ESO delivery. 

· The device is connected to the connectors defined in [AD 03]. 

Field Stabilization System interfaces:

The applicable documents for the FSS interface are [AD 03] and [AD 04]. 







Item
From
To
Signal
Description
Pin


COMMUNICATION:







1
LCU
ATS
STRAP board console
For maintenance only
RS232 Port
W31


APD HEAD CONTROL:







10
LCU
ATS
STRAP APD Head control lines
I/O lines to STRAP VME board
 


11
LCU
ATS
STRAP APD Head signal lines
See [AD 03] and [AD 04]
 


APD HIGH VOLTAGE:







20
LCU
ATS
High Voltage lines
See [AD 03] and [AD 04]
 


APD GATE CONTROL:







30
LCU
ATS
APD gate control +
contact closed to close the APD gate
Digital Output
 


31
LCU
ATS
APD gate control -
Common
 


32
 
 
 
 
 


33
LCU
ATS
APD gate control 50W coax
50W coax
 


34
 
 
 
 
 

<!-- /ESO-ICD -->
<!-- ESO-ICD -->
PRIVATE
E
S
O

O
n
l
y 
Field Stabilization System Software Interface:

The applicable document for the FSS software interface is [RD 07]. 







Item 
From 
To 
Signal 
Description 
Pin 


CONFIGURATION: 







2 
LCU 
ATS 
STRAP Static Configuration
(Can not be modified while APD is active) 
Static parameters:

· Open/Close loop 
· APD Gain 
· Diagnostic sampling period 
· Number of cycles 
· Hook-function name 
· etc ... (TBD) 
SW Command 
  

E
S
O

O
n
l
y
3 
LCU 
ATS 
STRAP Dynamic Configuration
(May be modified while APD is active) 
Dynamic parameters:

· Guide Reference Point 
· Interaction Matrix 
· Gate for APD gain control 
· etc ... (TBD) 
SW Command 
  


CONTROL: 







4 
LCU 
ATS 
Start STRAP 
start APD acquisition with current configuration
SW Command 
  


5 
LCU 
ATS 
Stop STRAP 
stop APD acquisition
SW Command 
  


6 
LCU 
ATS 
Gate APD gain 
turn gate on/off for APD gain control (used in chopping mode when alternating between object and sky)
SW Command or Digital Output 
  


7 
LCU 
ATS 
Update Interaction Matrix 
update interaction matrix
(for derotating M6 corrections)
SW Command 
  


8 
LCU 
ATS 
Request Interaction Matrix 
retrieve actual interaction matrix
SW Command 
  


9 
LCU 
ATS 
Update Reference Point
(see Note 1 below) 
update the APD reference point
SW Command 
  


STATUS & DIAGNOSTICS: 







10 
ATS 
LCU 
STRAP Status 
get STRAP status
SW Command / DB attribute 
  


11 
ATS 
LCU 
STRAP diagnostics parameters 
Diagnostics parameters:

· r0, t0, FWHM, SNR, centroid, extrema 
· etc ... (TBD) 
SW Command / DB attribute 
  


12 
LCU 
ATS 
Error Vector
(see Note 2 below) 
request the instantaneous Error Vector
SW Command 
  


SYNCHRONIZATION: 







13 
LCU 
LCU 
Function Hook
(see Note 3 below) 
hook a user-defined function synchronous to the APD control loop
API Function 
  


Note 1: The guide reference point is theoretically the optical center of the APD sensor. Since the FSS Translation Stage is a tracking axis, its position will be updated while tracking. However, it might be necessary to interpolate its position between 2 updates. This is achieved by offsetting the APD reference point synchronously with the Translation Stage control loop. The offset shall be less than 1/2 Airy disk (i.e. approx. 0.2" on the sky) so as not to loose the object.

Note 2: In order to compensate for tracking drift, the absolute offset of M6 is queried by the M6 Control Loop and sent to the telescope Alt/Az axes. The relative slow correction speed of the main axes allows not to recenter M6 but let the FS loop closed and active, that will bring M6 back to the center.

Note 3: The ATCS Requirements specify the need of synchronizing the acquisition system with the guiding loop. On the UTs, this is implemented between the TCCD system and the AG/FS loops by means of a user-defined function hook that is invoked at each exposure. The details of the function are application specific, the interface is TBD (shall try to match the existing one). The function shall have access to the actual parameters describing the data acquired during this cycle (error vector, SNR, FWHM etc...). 






<!-- /ESO-ICD -->
<!-- Standard footer -->

<!-- hhmts start -->Last modified: Wed Sep 20 15:24:28 METDST 2000 <!-- hhmts end -->




M6 drives

· Interface requirements for M6 electronics located on the fork at a distance <1.5m from M6 

· Electrical power: 230V UPS, wire cross section 0.75mm2 

· Cooling: one inlet outlet are reserved at the manifold distributor for M6 cabinet Thread type: BSP cylindrical 3/8". 

· Electrical signals: Terminal Block Y33 
· Optical signals: 2 ST Connectors 

· Housing of the M6 control electronics: Rittal cabinet 9U height mounted on the fork see Detector and M6 drives
<!-- ESO-ICD -->
PRIVATE

E
S
O

O
n
l
y


· Control of the High Frequency M6 tilt motions, required for field stabilization and chopping. For Adaptive Optics, a Deformable Mirror controlled by an appropriate electronics will replace M6. 

· The control is deployed on the Sensor LCU. 

· This sub-system belongs to the ESO delivery. 

· The device is connected to the Terminal Block Y33. 

M6 drives interfaces:

The connection layout and pin assignment of Y33 and ST1&ST2 is given in
Annex 1 VLT-DWG-ESO-15154-1787

 






Item 
From 
To 
Signal 
Description 


M6 TIP-TILT CONTROL: 






1 
LCU 
M6 
Tip tilt angles a &b 
16-bits command via optic fibre 


2 
 LCU 
M6 
Data Strobe + 
 COMACK Digital Output 


3 
 LCU 
M6 
Data Strobe Common 
Common 


4 
M6 
LCU 
X On-Target + 
XTARONI Input. Active high. Monitor Target-On on X channel Digital Input 


5 
M6 
LCU 
X On-Target Common 
Common 


6 
M6 
LCU 
Y On-Target + 
YTARONI Input. Active high. Monitor Target-On on Y channelDigital Input 


7 
M6 
LCU 
Y On-Target Common 
Common 


8 
 
 
 
 


M6 TIP-TILT MONITOR: 






20 
ATS 
LCU 
X-overflow + 
XOVERFI* Input. Active low. Overflow error indication on X channel Digital Input 


21 
ATS 
LCU 
X- overflow Common 
Common 


22 
ATS 
LCU 
Y- overflow + 
YOVERFI* Input. Active low. Overflow error indication on Y channel Digital Input 


23 
ATS 
LCU 
Y- overflow Common 
Common


24 
 
 
Power Monitor 
POWMONI Input Active high, Power monitor (M6 and Interface board)


25 
  
  
Shield 
  

<!-- /ESO-ICD -->
<!-- ESO-ICD -->
PRIVATE

E
S
O

O
n
l
y


Note: This device is controlled via the STRAP system. We need the following commands for centering and tilting M6. They are defined in the STRAP interfaces [RD 07].

M6 Software Interface:







Item
From
To
Signal
Description
Pin


CENTERING:







1
LCU
ATS
Center M6
Set M6 to Center Position M6
Software
Command
to STRAP


TILTING:







2
LCU
ATS
Set M6 absolute (a,b) tilt angles
Sets absolute M6 tilt angles (a,b) in radians
Software
Command
to STRAP


3
LCU
ATS
Set M6 relative (a,b) tilt angles
Sets relative M6 tilt angles (a,b) in radians



4
ATS
LCU
Get M6 absolute (a,b) tilt angles
Gets absolute M6 tilt angles (a,b) in radians


<!-- /ESO-ICD -->
<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 13:55:49 MET 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
FSS Filter Wheel

· Control of the position of the filter wheel that allows the insertions of 6 neutral density filters before the FSS. 
The motor and tacho are connected to the ESO standard VME4SA Servo Amplifier and the encoder to the MAC4 Motion Controller. The tacho generator is mounted on the motor. The velocity loop is closed in the amplifier. There is one switch used as a reference switch. 

· The control is deployed on the Sensor LCU. 

· The device is connected to the Terminal Block Z12. 

<!-- ESO-ICD -->
PRIVATE
E 
S 
O 
The motor control is a pure SW interface: the target position and Velocity are passed to the Motion Controller via the VME bus. The motor control software is provided by ESO.

<!-- /ESO-ICD -->
FSS Filter Wheel interfaces:

The connection layout and pin assignment is given in Annex 1 VLT-DWG-ESO-15154-1787

PRIVATE
Item 
From 
To 
Signal 
Description 

MOTOR: Minimotor DC Brush 2233-024S 
24VDC, 2.53W 





1
LCU
ATS
Motor M+
VME4SA Servo Amplifier Channel #1

2
LCU
ATS
Motor M-
VME4SA Servo Amplifier Channel #1

3
 
 
Shield MSH
 

TACHOGENERATOR: Minimotor 1.5 G 
1.5mV / rpm 





4
ATS
LCU
Tacho T+
VME4SA Servo Amplifier Channel #1

5
ATS
LCU
Tacho T-
VME4SA Servo Amplifier Channel #1

6
 
 
Shield TSH
 

ENCODER: Minimotor Optical HP HEDL 5540-500 
5 VDC - 500 counts/turn - Resolution = 7400 counts/degree = 0.5"/count 





7
LCU
ATS
Encoder 5V VCC
MAC4-INC Motion Controller Channel #1

8
LCU
ATS
Encoder GND
MAC4-INC Motion Controller Channel #1

10
ATS
LCU
Encoder A+
Differential line driver A+ 
MAC4-INC Motion Controller Channel #1

11
ATS
LCU
Encoder A-
Differential line driver A- 
MAC4-INC Motion Controller Channel #1

13
ATS
LCU
Encoder B+
Differential line driver B+ 
MAC4-INC Motion Controller Channel #1

14
ATS
LCU
Encoder B-
Differential line driver B- 
MAC4-INC Motion Controller Channel #1

16
ATS
LCU
Encoder I+
Differential line driver I+ 
MAC4-INC Motion Controller Channel #1

17
ATS
LCU
Encoder I-
Differential line driver I- 
MAC4-INC Motion Controller Channel #1

18
 
 
Shield 
 

SWITCHES: Micromat KS35A11 
24V, 4A max. - Repeatability=0.03degree 





19
ATS
LCU
Reference switch (NC)
MAC4-INC Motion Controller Channel #1

20
ATS
LCU
Reference switch(COM)
MAC4-INC Motion Controller Channel #1

21
 
 
Shield 
 

25
 
 
Global shield 
PE

 <!-- Standard footer --> 


<!-- hhmts start -->Last modified: Mon Dec 18 13:38:08 MET 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
FSS Aperture Field Diaphragm
<!-- ESO-ICD -->
  

PRIVATE
 
E
S
O

o  
n  
l  
y 

· Control of the Aperture Field Diaphragm before the FSS detector.  
The motor and tacho are connected to the ESO standard VME4SA Servo Amplifier and the encoder to the MAC4 Motion Controller. The tacho generator is mounted on the motor. The velocity loop is closed in the amplifier. There are two switches; one of them is used as a reference switch.  
  

· The control is deployed on the Sensor LCU.  

· The device is connected to the Terminal Block Y34.  

The motor control is a pure SW interface: the target reference position and Velocity are passed to the Motion Controller via the VME bus. The motor control software is provided by ESO.  

FSS Aperture Field Diaphragm interfaces:

The connection layout and pin assignment is given in Annex 1 VLT-DWG-ESO-15154-1787










Item 

From 

To 
Signal 

Description 



MOTOR: Model TBD 
TBDVDC, TBDW 










1

LCU

ATS
Motor M+

VME4SA Servo Amplifier Channel #2



2

LCU

ATS
Motor M-

VME4SA Servo Amplifier Channel #2



3

 

 
Shield MSH

 



TACHOGENERATOR: Model TBD 
TBDmV / rpm 










4

ATS

LCU
Tacho T+

VME4SA Servo Amplifier Channel #2



5

ATS

LCU
Tacho T-

VME4SA Servo Amplifier Channel #2



6

 

 
Shield TSH

 



ENCODER: Model TBD 
5 VDC - TDB counts/turn - Resolution = TBD counts/degree = TBD"/count 










7

LCU

ATS
Encoder 5V VCC

MAC4-INC Motion Controller Channel #2



8

LCU

ATS
Encoder GND

MAC4-INC Motion Controller Channel #2



10

ATS

LCU
Encoder A+

Differential line driver A+ 
MAC4-INC Motion Controller Channel #2



11

ATS

LCU
Encoder A-

Differential line driver A- 
MAC4-INC Motion Controller Channel #2



13

ATS

LCU
Encoder B+

Differential line driver B+ 
MAC4-INC Motion Controller Channel #2



14

ATS

LCU
Encoder B-

Differential line driver B- 
MAC4-INC Motion Controller Channel #2



16

ATS

LCU
Encoder I+

Differential line driver I+ 
MAC4-INC Motion Controller Channel #2


E
S
O

o  
n  
l  
y 

17

ATS

LCU
Encoder I-

Differential line driver I- 
MAC4-INC Motion Controller Channel #2



18

 

 
Shield 

 



<div align="center">SWITCHES: Model TBD 
TBDV, TBDA max. - Repeatability=TBDdegree </div>










18

ATS

LCU
Reference switch (NC) 

 MAC4-INC Motion Controller Channel #2



18

ATS

LCU
Reference switch (COM) 

 MAC4-INC Motion Controller Channel #2



18

 

 
Shield 

 



18

 

 
Global shield 

 PE

PRIVATE


<!-- /ESO-ICD -->
<!-- Standard footer -->

<!-- hhmts start -->Last modified: <!-- #BeginDate format:aEn2m -->Wed 16-May-2001 17:26<!-- #EndDate --><!-- hhmts end --> 




<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
FSS Translation Stage

· Control of the XY translation stage used as support for the FSS Detector, including motor, tachometer, encoder and reference switches. 
The motor and tacho are connected to the ESO standard VME4SA Servo Amplifier and the encoder to the MAC4 Motion Controller. The tacho generator is mounted on the motor. The velocity loop is closed in the amplifier. There are two switches per axis: a positive limit switch (PLS) and a negative limit switch (NLS). One of them is used as reference switch. 

· The control is deployed on the Sensor LCU. 

· The device is connected to the Terminal Blocks Z16 for the X-axis and Z17 for the Y-axis. 

<!-- ESO-ICD -->
  

PRIVATE
E 
S 
O 
The motor control is a pure SW interface: the target position and velocity are passed to the Motion Controller via the VME bus. The motor control software is provided by ESO. 

<!-- /ESO-ICD -->
FSS Translation Stage interfaces:

The connection layout and pin assignment is given in Annex 1 VLT-DWG-ESO-15154-1787

PRIVATE
X-Axis 





Item 
From 
To 
Signal 
Description 

MOTOR: Minimotor DC Brush 2233-024S 
24VDC, 2.53W - GearBox = 1:29.6 - Screw Pitch = 1mm 





1
LCU
ATS
Motor M+
VME4SA Servo Amplifier Channel #3

2
LCU
ATS
Motor M-
VME4SA Servo Amplifier Channel #3

3
 
 
Shield MSH
 

TACHOGENERATOR: Minimotor 1.5 G 
1.5mV / rpm 





4
ATS
LCU
Tacho T+
VME4SA Servo Amplifier Channel #3

5
ATS
LCU
Tacho T-
VME4SA Servo Amplifier Channel #3

6
 
 
Shield TSH
 

ENCODER: Minimotor Optical HP HEDL 5540-500 
5 VDC - 500 counts/turn - Resolution in Quadrature = 59200 counts/mm = 16.9nm/count 





7
LCU
ATS
Encoder 5V VCC
MAC4-INC Motion Controller Channel #3

8
LCU
ATS
Encoder GND
MAC4-INC Motion Controller Channel #3

10
ATS
LCU
Encoder A+
Differential line driver A+ 
MAC4-INC Motion Controller Channel #3

11
ATS
LCU
Encoder A-
Differential line driver A- 
MAC4-INC Motion Controller Channel #3

13
ATS
LCU
Encoder B+
Differential line driver B+ 
MAC4-INC Motion Controller Channel #3

14
ATS
LCU
Encoder B-
Differential line driver B- 
MAC4-INC Motion Controller Channel #3

16
ATS
LCU
Encoder I+
Differential line driver Z+ 
MAC4-INC Motion Controller Channel #3

17
ATS
LCU
Encoder I-
Differential line driver Z- 
MAC4-INC Motion Controller Channel #3

18
 
 
Shield
 

SWITCHES: Micromat KS35A11 
24VDC, 4A max- Repeatability=0.02mm 





19
ATS
LCU
Negative Limit (NC)
MAC4-INC Motion Controller Channel #3

20
ATS
LCU
Negative Limit (COM)
MAC4-INC Motion Controller Channel #3

21
 
 
Shield 
 

22
ATS
LCU
Positive Limit (NC)
MAC4-INC Motion Controller Channel #3

23
ATS
LCU
Positive Limit (COM)
MAC4-INC Motion Controller Channel #3

24
 
 
Shield
 


  

PRIVATE
Y-Axis 





Item 
From 
To 
Signal 
Description 

MOTOR: Minimotor DC Brush 2233-024S 
24VDC, 2.53W - GearBox = 1:29.6 - Screw Pitch = 1mm 





1
LCU
ATS
Motor M+
VME4SA Servo Amplifier Channel #4

2
LCU
ATS
Motor M-
VME4SA Servo Amplifier Channel #4

3
 
 
Shield MSH
 

TACHOGENERATOR: Minimotor 1.5 G 
1.5mV / rpm 





4
ATS
LCU
Tacho T+
VME4SA Servo Amplifier Channel #4

5
ATS
LCU
Tacho T-
VME4SA Servo Amplifier Channel #4

6
 
 
Shield TSH
 

ENCODER: Minimotor Optical HP HEDL 5540-500 
5 VDC - 500 counts/turn - Resolution in Quadrature = 59200 counts/mm = 16.9nm/count 





7
LCU
ATS
Encoder 5V VCC
MAC4-INC Motion Controller Channel #4

8
LCU
ATS
Encoder GND
MAC4-INC Motion Controller Channel #4

10
ATS
LCU
Encoder A+
Differential line driver A+ 
MAC4-INC Motion Controller Channel #4

11
ATS
LCU
Encoder A-
Differential line driver A- 
MAC4-INC Motion Controller Channel #4

13
ATS
LCU
Encoder B+
Differential line driver B+ 
MAC4-INC Motion Controller Channel #4

14
ATS
LCU
Encoder B-
Differential line driver B- 
MAC4-INC Motion Controller Channel #4

16
ATS
LCU
Encoder I +
Differential line driver I+ 
MAC4-INC Motion Controller Channel #4

17
ATS
LCU
Encoder I-
Differential line driver I- 
MAC4-INC Motion Controller Channel #4

18
 
 
Shield 
 

SWITCHES: Micromat KS35A11 
24VDC, 4A max- Repeatability=0.02mm 





19
ATS
LCU
Negative Limit (NC)
MAC4-INC Motion Controller Channel #4

20
ATS
LCU
Negative Limit(COM)
MAC4-INC Motion Controller Channel #4

21
 
 
Shield
 

22
ATS
LCU
Positive Limit (NC)
MAC4-INC Motion Controller Channel #4

23
ATS
LCU
Positive Limit (COM)
MAC4-INC Motion Controller Channel #4

25
 
 
Global shield 
PE

<div align=right><!-- Standard footer --></div>

<!-- hhmts start -->Last modified: Mon Dec 18 13:41:09 MET 2000 <!-- hhmts end -->



M6 Deformable Mirror
<!-- ESO-ICD -->
PRIVATE

E
S
O

O
n
l
y

 

This function will not be implemented in the initial phase of the project, but at a later stage. However interface requirements related to reserved space, cable routing, plug design, etc... apply. 

 

· Control of the M6 Deformable Mirror for Adaptive Optics. 


· The control is deployed on the Adaptive Optics LCU. 

· This sub-system belongs to the ESO delivery. 

<!-- /ESO-ICD -->
<!-- Standard footer -->

<!-- hhmts start -->Last modified: Wed Aug 4 07:54:19 METDST 1999 <!-- hhmts end -->



Auxiliary LCU

LCU Configuration

<!-- ESO-ICD -->
PRIVATE
E
S
O
PRIVATE
Board
Supplier
Number



CPU
Motorola MVME 167 (8MB)
1



Digital I/O Interface 
Acromag AVME 9481 
1 



Serial Interface
ESD ISER8
1


<!-- /ESO-ICD -->
This LCU is responsible for:

PRIVATE
Device 
Connector 

Enclosure
Anemometers 
W42


Air Conditioning
M1 Temperatures 
W41
W41 

Relay Optics Structure (ROS) Shutter 
W42 

Lighting 
<div align="center">Z18 </div>

Transporter 
W42 

Service modules: 

· Auxiliary power 

· Hydraulic and Pneumatic systems 

· Liquid Cooling Module 
W42
W42
W41 

<!-- /ESO-ICD -->
Schematic:

PRIVATE "TYPE=PICT;ALT=[Auxiliary LCU]"

<!-- Standard footer -->

<!-- hhmts start -->Last modified: <!-- #BeginDate format:Sw1m -->14 May, 2001 11:00<!-- #EndDate --><!-- hhmts end -->



Enclosure and Anemometers

· The telescope enclosure sub-system consists of the 2 half-shells of the enclosure slit, the associated seals, the six (6) clamping actuators and 2 telescope anemometers 3D, located on top of the enclosure and mounted 45 degrees apart the local vertical. This configuration allows proper computation of the wind speed and direction independently of the opening angle of the enclosure shells.
The slit open/close motion is performed to close the enclosure of the telescope during day-time and so to be able to protect the equipment and control the temperature inside the telescope in the building. When the wind speed, measured on the local anemometers, exceeds 18 m/s, the opening of the slit is not allowed; if the slit is open, it is automatically closed if possible.
The Remote control of the enclosure sub-system is only possible when the corresponding button of the enclosure control panel is in REMOTE position.
All signals described in the table below are mapped into the PLC and are accessible via status commands. 

· The control is deployed on the Auxiliary LCU. 

· The Enclosure is controlled by a Siemens PLC, connected to a RS422 Serial Port via Connector W42. 

· The 2 anemometers 3D are connected to the Siemens PLC. Their monitoring is performed via the Serial link. 

Enclosure interfaces:

PRIVATE
Item 
From 
To 
Signal 
Description 

COMMUNICATION: 





1 
  
  
Enclosure and Anemometer Communication 
RS422 Port to Siemens PLC
Protocol RK512/3964R 

 

PRIVATE
Item 
From 
To 
Signal 
Description 

SEALS STATUS: 





10 
ATS 
LCU 
Enclosure seals inflated + shells clamped status 
Bit set if shells clamped and seals fully inflated
PLC Digital Input 

11 
ATS 
LCU 
Enclosure seals deflated + shells unclamped status 
Bit set if seals fully deflated and shells unclamped
PLC Digital Input 

GLOBAL STATUS: 





20 
ATS 
LCU 
Enclosure closing automatically 
contact closed if Enclosure is closing automatically
PLC Digital Input 

21 
ATS 
LCU 
Enclosure local status 
contact closed if Enclosure is in local control
PLC Digital Input 

ANEMOMETER MONITOR: 





30 
ATS 
LCU 
Anemometer #1
wind speed in X ,Y,Z 
Software command to W42
X,Y,Z coded over 8 Bits in m/s
Note: directions are related to the Anemometer, hence relative to the opening of the Enclosure shells. 

32 
ATS 
LCU 
Anemometer #2
wind speed in X,Y,Z 








 

PRIVATE
Item 
From 
To 
Signal 
Description 

ENCLOSURE CONTROL: 





100 
LCU 
ATS 
Enclosure front shell open command 
asserted to open the enclosure front shell at the next start command.
negate to close at the next start command. 

101 
LCU 
ATS 
Enclosure rear shell open command 
asserted to open the enclosure rear shell at the next start command.
negate to close at the next start command. 

102 
LCU 
ATS 
Enclosure front shell start command 
assert to initiate the opening/closing of the front shell. Negate to stop the motion whenever the desired postion is reached. 

103 
LCU 
ATS 
Enclosure rear shell start command 
assert to initiate the opening/closing of the rear shell. Negate to stop the motion whenever the desired postion is reached. 

104 
ATS 
LCU 
Enclosure front shell actuator#1 
  
  Software command to W42
Actuator position coded over 8 Bits
Scaling t.b.d. 

  

  

105 
ATS 
LCU 
Enclosure front shell actuator#2 


106 
ATS 
LCU 
Enclosure rear shell actuator#3 


107 
ATS 
LCU 
Enclosure rear shell actuator#4 


<!-- Standard footer -->

<!-- hhmts start -->Last modified: <!-- #BeginDate format:aEn2m -->Fri 25-May-2001 11:00<!-- #EndDate --><!-- hhmts end -->




<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Air Conditioning

· The air conditioning control is under the contractor's responsibility. It is controlled via a RS232 serial interface and provides the commands to switch on/off the Air Conditioning, read the enclosure temperature and set the reference temperature. 
  

· The control is deployed on the Auxiliary LCU. 

· The Air Conditioning is controlled by a Sauter PLC, connected to a RS232 Serial Port via Connector W41. 

· The temperature values of the M1 sensors are input to the Sauter PLC, via the Connector W41. 

· The Air Conditioning Control System gives a Warning flag when the Dew point temperature is reached. The Dew point sensor readout is available to ESO via the RS232. The parameter Dew point Margin Tdm can be adjusted via the RS232. 

Air Conditioning interfaces:

The connection layout and pin assignment is given in Annex 1 VLT-DWG-ESO-15154-1769 
  

PRIVATE
Item
From
To
Signal
Description

COMMUNICATION:





1
 
 
Air Conditioning Communication
RS232 Port to Sauter PLC 
Protocol details in [RD 10]

AIR CONDITIONING:





10
ATS
LCU
Air Conditioning status
Get status information on the air Conditioning system

11
ATS
LCU
Cooling water temperature
Get cooling water temperature

12
ATS
LCU
Enclosure temperature
Get enclosure temperature

13
LCU
ATS
Start/Stop Air Conditioning
· Start 
using the reference temperature 

· Stop  

14
LCU
ATS
Air Conditioning Reference Temperature
Set the Reference Temperature

15
LCU
ATS
Dew point sensor 
Get Dew point sensor value 

16
LCU
ATS
Dew point sensor Temperature Margin
Set the Dewpoint Temperature Margin

17
LCU
ATS
Dew Point Warning
warning Flag ON when Dew Point is reached inside the enclosure during Daytime Stand By mode.

M1 TEMPERATURES:





20
LCU
ATS
Temperature M1 A,B,C
Mirror M1 temperatures as read from Altitude LCU


  

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 10:14:16 MET 2000 <!-- hhmts end -->



Relay Optics Structure (ROS) Shutter

· Control of the Relay Optics Structure Shutter 


· The control is deployed on the Auxiliary LCU. 

· The device is connected to a RS422 Serial Port via Connector W42 (TBC). 

ROS Shutter interfaces:

PRIVATE
Item 
From 
To 
Signal 
Description 
Pin 

COMMUNICATION: 






1
 
  
Communication 
RS422 Port 
W42 

CONTROL: 




Software
Command
to W42 

10
LCU
ATS 
ROS Shutter Close command 
Bit set to close the ROS Shutter 


11
LCU
ATS 
ROS Shutter Open command 
Bit set to open the ROS Shutter 


STATUS: 






20
ATS
LCU 
ROS Shutter Close Status 
Bit set if ROS Shutter is fully closed 


21
ATS
LCU 
ROS Shutter Open Status 
Bit set if ROS Shutter is fully open 


 

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Wed Sep 20 15:24:25 METDST 2000 <!-- hhmts end -->



Transporter

· The transporter is the mechanical part of the AT that is involved in the relocation procedure and hosts the anchoring system. 

· The control is deployed on the Auxiliary LCU. 

· The devices are controlled by the Siemens PLC, connected to a RS422 Serial Port via Connector W42. 

There will be 10 relative simple on-off drives connected to a manual control panel, for the following functions:

1. telescope engage/disengage

2. transporter engage/disengage

3. telescope up and down

4. transporter wheels up and down

5. transporter direction

6. transporter forwards and backwards

7. ROS up and down

8. enclosure open and close

9. ROS lock and unlock

The complete control of the transporter drives for the whole relocation sequence is under the contractor's responsibility, but status information is provided to the ESO control system via signals connected to the LCUs Digital I/O board.

Transporter interfaces: 

PRIVATE
Item
From
To
Signal
Description

COMMUNICATION: 





1
 
 
Transporter Communication
RS422 Port to Siemens PLC - 
Protocol RK512/3964R

TRANSPORTER:





10
ATS
LCU
Transporter status
contact closed if transporter is operational,
i.e. all following modules are in Normal situation: 

· Oil level and Temperature switch of the hydraulic module 

· Compressed air OK (t.b.c.) 

· Battery & Electrogen group 

11
ATS
LCU
Relocation status
Double contact:
one contact to altitude & one contact to azimuth: 
Contacts closed if relocation is completed,
i.e. all following conditions are verified: 

· Transporter engaged 

· Telescope engaged 

· ROS engaged against Telescope 

· 400V-OK 

· Electro-Group command OFF 

· Hydraulic module switched off 

· Pneumatic module switched off 

· All panels in Off or Remote position
(fixed control, enclosure, air conditioning & LCM panels) 

12
ATS
LCU
Motion Stop status
Double contact:
one contact to altitude & one contact to azimuth: 
Contact open if a Motion Stop is activated,
i.e. the following Motion stops are activated: 
the motion stop identification is done via the serial port 

· List of Motion stops 

 

<!-- Standard footer -->

<!-- hhmts start -->Last modified: <!-- #BeginDate format:aEn2m -->Fri 25-May-2001 11:31<!-- #EndDate --><!-- hhmts end -->



Service modules

They control and/or monitor the status of: 

· Auxiliary power
The status of the Auxiliary Power System shall be interfaced to the digital I/O interface. 
(See drawing Power distribution schematics in section 3.1.1 on page 8). 

· The Hydraulic and Pneumatic systems (HPS) provide hydraulic and pneumatic power to the transporter and the enclosure.

· The Liquid Cooling module (LCM) is providing cooling power to the Air Conditioning, the Thermal Control of the electro-mechanical parts and to the FAS and FSS detectors. 

· The control is deployed on the Auxiliary LCU. 

· The Power, Hydraulic and Pneumatic modules are controlled by the Siemens PLC, connected to a RS422 Serial Port via Connector W42. 

· The Liquid Cooling module is controlled by the Sauter PLC, connected to a RS232 Serial Port via Connector W41. 

Service modules interfaces:

PRIVATE
Item

From
To
Signal
Description
Pin

COMMUNICATION: 







1 

ATS 
LCU 
Power, Hydraulic & Pneumatic Systems 
RS422 Port to Siemens PLC - Protocol RK512/3964R 
W42 

2 

ATS 
LCU 
Liquid Cooling Module 
RS232 Port to Sauter PLC - Protocol details in [RD 10] TBD 
W41 

 

 







E
S
O
Item
From
To
Signal
Description
Pin


<div align="center">MONITORING: </div>







3 
  
  
Flow Meter monitoring 
Digital I/O 
Y35 

 

PRIVATE
Item

From
To
Signal
Description
Pin

AUXILIARY POWER: 







10 

ATS 
LCU 
main power status 
contact open if undervoltages on power line from station plug 
Software Command to W42

11 

ATS 
LCU 
Battery status 
contact open if battery is NOT OK
Digital Input 


HYDRAULIC & PNEUMATIC SYSTEMS: 







20 

LCU 
ATS 
Hydraulic System control 
Switch On/Off if control panel button is in Remote position 
Software Command to W42

21 

ATS 
LCU 
Hydraulic System status 
  


22 

LCU 
ATS 
Pneumatic System control 
Switch On/Off if control panel button is in Remote position 


23 

ATS 
LCU 
Pneumatic System status 
  


LIQUID COOLING MODULE: 







30 

ATS 
LCU 
LCM Status 
Complete LCM status 
and Liquid cooling Temperature 
Software Command to W41

E
S
O
31
LCU
ATS
Flow-meter +
+24VDC, 
Y35_1


32
ATS
LCU
Flow-meter signal
Frequency as a function of flow
Y35_2


33
 
 
Shield (Y35_1 - Y35_2)
 
Y35_3

 

<!-- Standard footer -->


<!-- hhmts start -->Last modified: <!-- #BeginDate format:aEn2m -->Fri 25-May-2001 11:05<!-- #EndDate --><!-- hhmts end -->




Interlocks

The Interlocks are dependent on the design. They are to be defined during design by the contractor, according to the safety requirements, manual operation and maintenance needs. 

ESO has defined three types of interlocks: 

· Motion Stop Interlocks: general interlocks generated by actuation of the Motion Stop switches mounted in the AT. One of the push-buttons must be of type locking key. All Motion Stop have the same action; their location and quantity are defined in AMOS Control System document [RD03], Chapter 12. 

· Local Interlocks: these interlocks are used to prevent specific actions between equipment. 

· Equipment Interlocks: these interlocks are used to prevent specific actions within an equipment. 

The hazard analysis is provided in [AD06]. 

Interlocks Table:

PRIVATE
Subsystem

Action
Cause
Recovery Action

Generating
Receiving




VLTI Motion Stop
in Control Room
All drives within AT
Main Drives immobilized
Except small axes (NFD, CFD, FSS, M2, M10).
Button pressed
or key removed
Release button
or Replace key

AT Motion Stop
on AT





Hardware Limit
Azimuth
Drive immobilized
Limit Switch
Move back into stroke

Hardware Limit
Altitude
Drive immobilized
Limit Switch
Move back into stroke

Interior Door
All drives
Drives immobilized
Limit Switch
Close the door

Alt-Locking pins
Altitude
Drive immobilized
Limit Switch
Put Alt-Locking pins into unlocking position

Transporter PLC
All subsystems except those supplied by battery and small axes
Power Shutdown
Watchdog
Reset Transporter PLC

Altitude Parking
Telescope Up/Down
Positive direction inhibited
Limit Switch
Put Altitude axis in Parking position

Stroke limit
Azimuth & Azimuth Cable Wrap
Drive immobilized
Azimuth CWP Limit Switches
Move back Az axis into stroke

Altitude brakes
Altitude
Drive immobilized
Limit Switch
Release brakes

Azimuth brakes
Azimuth
Drive immobilized
Limit Switch
Release brakes

Enclosure Area
Azimuth + Enclosure
Drives immobilized
Key removed
Replace key

Transporter PLC
Enclosure
Close Winshield
Anemometer give too high windspeed
none 

Station Plug
Transporter
All transporter drives immobilized
Plug not in Travel position
Unplug AT
or plug in Travel Position

 

Important Note: Before undertaking any recovery action, one has to have well understood why the corresponding interlock has been activated and take the corrective action to prevent from activating the interlock in the future. 

Moreover, there are many internal locks. The Transporter PLC checks the feasibility of each operator command, the Sauter PLC is equipped with a watchdog that stops the air conditioning system and switches the alarm status on, each electrical motor is equipped with PTC sensors, each inverter is equipped with overcurrent protection. See [RD 03] and [AD 06]. 

 

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 10:22:27 MET 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Electrical & Optical Interfaces

Station Plugs

The station plugs accommodate the cables and optical fibres provided by ESO at every station; the details of the implementation is given in Station Plug Interface[RD 14] 

The main characteristics of the station plug are: 

· 400V Non-UPS electrical power + shield - 

· 12 Optical fibres - multi-mode graded index with core diameter of 62.5 mm and cladding F125mm. 
The optical fibres will carry: 

· Control LAN 

· Guide/Acquisition LAN 

· Time Bus 

· High data rate links 1 & 2. 

· 24 Electrical lines Section 0.14mm2 + Shield to PE 

· 2 for Motion Stop 

· 8 for station ID 

· 1 for carriage interlock 

· +24VDC line 

· 12 spare lines 

Cables

ESO provides the guiding and mechanical fixation of the cables connected between LCU and terminal blocks (in order to avoid mechanical stresses on electrical connections). All other cables as described in Installation of ESO cables [RD 17] shall be routed and mechanically fixed by the contractor, except Nasmyth and Pupil beacon fiber. 

 

  

The cables labelling is described in [RD 17] It follows the template Ca ESO/b /c where:
a is a sequential number 
b is the input connection location 
c is the output connection location 

The locations and their abbreviation is described in annex 3 of [RD 17] and in drawing Telescope control cabling path of AN01 

  

PLC I/O

PRIVATE
I/O characteristics description
Input
Output
Remark

Siemens PLC on Transporter

PNP Open collector 
Load between output and ground

PHASE Drive controller

NPN Totem pole
Load between output and ground

<!-- Standard footer -->

<!-- hhmts start -->Last modified: Mon Dec 18 10:22:04 MET 2000 <!-- hhmts end -->



<!doctype html public "-//w3c//dtd html 4.0 transitional//en">
Mechanical Interfaces

Local Control Units


The table sumarizes the volume and location requirements for ESO LCUs and network equipment. 
  

PRIVATE
Equipment
Size
Cabinet
Location

Altitude LCU
19" VME Chassis 
Depth: 500mm 
Height: 7U 
(6 for VME chassis + 1 for fan)
Altitude 
See [RD 03], Section 5.1
Fork B side 

Azimuth LCU

The entire left part of the Signal Cabinet is let for integration of ESO equipment. 
See [RD 03], Section 5.4
Transporter

Sensor LCU




Auxiliary LCU




Adaptive Optics LCU




Dual Feed LCU




Network Equipment
3U x 500mm



Access to the back of each LCU cabinet is provided.
The signal cabinet will be movable for access to the rear side and to the GIS signal connection point
The Altitude cabinet will be mounted on sliding rails for easier maintenance access. 

  

Detector and M6 drives

The mechanical interfaces of the Field Acquisition Sensor (TCCD + ACE controller) and the Field Stabilization Sensor (STRAP APD) are defined in AD 09, AD 10 and Drawing 03 of the Technical Specifications [RD 02]. 

Space and connection means shall be provided from Sensor LCU to FAS via the ROS plug for fibre installation (see [RD 02],  section 5.1.14). 

Space and cooling capacity for electronics shall be provided in the vicinity (< 1.5 m) of M6. The minimum space required is: Width:19" VME Chassis, Height: 8HU, Depth: 600 mm including 100mm for connectors. 

Pupil Beacon

Space and connection means shall be provided from Altitude LCU to M2 via the tube cable wrap for Pupil beacon fibre installation.see [RD15] 

Nasmyth focus

In view of future installation of specific ESO equipment, the following requirements apply: 

Nasmyth Beacon: 
Space and connection means shall be provided from Altitude LCU to Nasmyth Wheel for Nasmyth beacon fibre installation (see [RD 02],  section 5.1.11). 

Future Extensions

This section defines the internal interface requirements applicable to the ATS design in order to allow future installation of upgrades by ESO. 

Adaptive Optics

Coudé focus environment

In view of the future installation of the Wave Front Sensor for adaptive optics, the following interface requirements apply: 

Space and connection means for cable installation between the Wave Front Sensor and the Terminal Block located inside the Signal Cabinet shall also be provided. 
Requirements: 
Flanges shall be reserved for later installation of an additional Transporter/RO plug to accomodate the cables needed for future Adaptive Optics and Dual Feed. This additional plug will be installed by ESO. Its supply is outside the contract's scope. This plug shall be able to accommodate the most constraining of the two following connection schemes: 

· A) 100 analogue signals pins + 2 fibre connections 

· B) 15 analogue signal pins + 28 fibre connections 

Space for Cables: 
The space to be reserved for cable routing shall accommodate a cable fitting a 20mm thick rectangle (equivalent section of 30mm diameter) with a minimum bending radius of R=100mm running from the Transporter/RO plug to the two corresponding reserved spaces DWG7 and DWG8 of [RD 02]. 

Adaptive Optics LCU

Dual Feed

Coudé focus environment

In view of the future installation of the Dual Feed system at the Coudé focus, the following interface requirements apply: 

Space and connection means for cable installation between M10df & M11b and the Terminal Block located inside the Signal Cabinet shall be provided, 
Requirements are already included in those defined in previous section (cables for AO and WFS) 

TBC: Space and cooling capacity for electronics shall be provided in the vicinity (< 1.5 m) of M10df and of M11b. The minimum total space required is 300 x 150 x 250 mm3. 

Dual Feed LCU

The space required is defined in Local Control Units. 

Atmospheric Dispersion Corrector

The Relay Optics Structure shall provide an interface for future installation of an Atmospheric Dispersion Compensator (ADC). The following requirements apply: 

Space and connection means for cable installation between the ADC and the Terminal Block located inside the Signal Cabinet shall be provided with the following characteristics: 

· cable diameter: 20mm, 

· bending radius: 100mm, 

· Connector: 25poles, 1A, low-voltage. 

Connection type & location

Terminal blocks

They are used for common electrical connections; these are referenced Zn and Yn. 
Terminal block type:  (eg. Phoenix MTTB1.5, or Wago TBC) 
Industrial Type 
WAGO INTERFACE MODULE No.289-533 (Z1, Z5)
WAGO INTERFACE MODULE No.289-556 (Y33)
WAGO INTERFACE MODULE No.289-546 (Y333)
WAGO X-COM Serie 870 No.870-101( Z11, Z21, Z22, Z23)

ESO custom made Type:    
ATS-MODULE VME4SA-1 INTERFACE (Z3, Z10, Z12, Z16, Z17, Y20, Z34)
ESO-I/O Interface Adapter 9916-21 (Z18)

Serial interfaces

They will be connected through standardized DB9 connectors grouped together on a panel with front access. These are referenced Wn. 

Heidenhain encoders

They will connect directly to the LCU, front panel of the IK320 interface. Connector type DB9 

Optical fibre connection

See also [RD17] installation of ESO cables 

· Station Plug connector type: Harting HAN Modular 

· OpticalLAN connector type: ST 

· TIM connector type: ST 

· HDRL connector type: ST 

· Sensor LCU to Field Acquisition Sensor (TCCD) connector type: SMA 

· Pupil Beacon connector type: TBD 

· Nasmyth Beacon connector type: TBD 

ESO Terminal Blocks Yn
Terminal blocks referenced Yn are part of ESO supply; space shall be provided by the contractor, preferably in the vicinity of the Zn terminal blocks. 
The same mounting supports as for Zn terminal blocks is required. Anticipated Terminal Block Yn length: 20cm (TBC). 

  

Cooling

The cooling of the racks is under the contractor's responsibility. See [RD 03], Chapter 5. 

<!-- Standard footer -->

<!-- hhmts start -->Last modified: <!-- #BeginDate format:En2 -->23-May-2001<!-- #EndDate --><!-- hhmts end --> 




____oOo____

