## Slide 1: NGCII System and Hardware
Mathias Richerzhagen (Speaker)
Matthias Seidel
Alexander Rüde
Omar Sqalli

## Slide 2: Introduction
Today, we will answer the FAQ for NGCII
What does it do?
What’s in the box?
How do I use it?
When can I have one?
?
?
?
?

## Slide 3: Background Information

## Slide 4: Background Information
Published journal article in Journal of Astronomical Telescopes, Instruments, and Systems VOL. 10 · NO. 4 | October 2024
  https://doi.org/10.1117/1.JATIS.10.4.046001
  Preprint available at:
  https://doi.org/10.48550/arXiv.2410.22965
Published Information on NGCII

## Slide 5: Background Information
Available ICDs
https://pdm.eso.org/kronodoc/HQ/N/473529

## Slide 6: What does it do?

## Slide 7: What does it do?
NGCII is the detector controller for ELT and new VLT instruments
Manufacturer | Detector | Type
Teledyne | HAWAII 2RG | CMOS
Teledyne | HAWAII 4RG | CMOS
Leonardo | SAPHIRA | eAPD CMOS
Teledyne E2V | CCD231-84 | CCD
Teledyne E2V | CCD290-99 | CCD
Teledyne Imaging Sensors | GEOSNAP | Fully Digital ROIC
Manufacturer | Detector | Type
Leonardo | Large SAPHIRA | eAPD CMOS
LBNL | LBNL CCD | Fully Depleted CCD
SOFRADIR | ALFA | CMOS
Raytheon | AQUARIUS | -
Raytheon | VIRGO | -
STA | STA CCD | -
Gpixel | GSENSE | -
Supported Detectors
Roadmap

## Slide 8: What does it do?
Focus on operational stability
NGC1 can read listed detectors limited by detector noise, but…
External Interfaces not ELT compliant
Combined CCD/CMOS circuits
limit performance of both
Limited readout speed
All interfaces Ethernet based
Dedicated circuits for CCD, CMOS and fully digital detectors
Greatly improved internal bandwidth

## Slide 9: What does it do?
Working well for CMOS already
HAWAII-2RG
Engineering Grade
Detector

## Slide 10: What does it do?
Working well for CMOS already
HAWAII-2RG fast mode noise is detector limited as for NGC1.
HAWAII-2RG slow mode achieving noise levels not seen before at <10e− for simple CDS.

## Slide 11: What does it do?
Working well for CMOS already
HAWAII-2RG running at 30fps without observable hot-pixel trailing

## Slide 12: What is in the box?

## Slide 13: What is in the box?
Parts of the NGCII system

## Slide 14: What is in the box?
Parts of the NGCII system
Included:
NGCII Detector Control Hardware
NGCII Detector Workstation
NGCII Firmware
NGCII Software
  NGCII/Core
  NGCII/IR
  NGCII/OPT
Not Included:
Communication Infrastructure
Detector
Preamplifier
Cables
Some of those may be part of the detector work-package but not NGCII so still a deliverable of ESO.

## Slide 15: Detector Workstation

## Slide 16: What is in the box?
  Server PC selected from the list of ESO computer and network standards.
  NGCII makes sure that at least one current configuration is approved for use as DWS.
  Runs ELT Software Environment
  Currently: Dell R7525
    AMD Epyc CPU
    128GB RAM
    Intel X710-DA2 NIC
Parts of the NGCII system – Detector Workstation
Some selection criteria:
Highest single core performance available
Known-working network controller
Core affinity of network controller

## Slide 17: Detector Control Hardware

## Slide 18: What is in the box?
  2HU 19” rack mount modular card carrier.
  Based on MicroTCA.4
  Controls 1 or 2 detectors. E.g.:
    2x HAWAII-2RG with 32 outputs
    2x CCD231
    1x HAWAII-4RG with 64 outputs
    1x CCD290
Parts of the NGCII system – Detector Control Hardware

## Slide 19: What is in the box?
De-facto standard for modular systems in the large science community
  Industrial standard hosted by PICMG
    ESO is an associate member
  Annual conference
  Paid support / consulting available
  Solid industry base

## Slide 20
Advanced
Mezzanine
Cards [AMC]
Rear-Transition
Modules [RTM]
Power Supply Unit [PSU]
MicroTCA
Carrier Hub [MCH]
Zone 3
Zone 1
Zone 1 Backplane
PCIe Gen3 x8
12V Power
Management
Etc..
Control and Data
Ethernet
Power
Detector
Interface

## Slide 21: What is in the box?
Example System: METIS SCAO, SAPHIRA32
NGCII – ICD and Architecture Specification – SAPHIRA
https://pdm.eso.org/kronodoc/HQ/ESO-410219
for other system configurations see respective architecture specification…
Shelf + MCH + PSU
IOxOS
IFC1414
2x ESO
AQ22 AMC
AIES
MFMC
2x CMOS AQ22 RTM
APD Bias
RTM
CMOS C20B20 RTM

## Slide 22: NGCII Firmware

## Slide 23: What is in the box?
Parts of the NGCII system – Firmware
Concept approved by external audit!

## Slide 24: NGCII Software
See the following talk by Jörg Stegmeier!

## Slide 25: How to use it?

## Slide 26: Communication Interfaces

## Slide 27: How to use it? :: Communication
Communication Interfaces
  Detector Workstation
    Multiple network interfaces.
    Exact rules for connection described in ESO-566602
  Detector Control Hardware
    Four SFP+ cages on the front panel of the host controller module.
    Exact rules for connection described in ESO-566603
Easier rules of thumb on following slides!

## Slide 28: How to use it? :: Communication
Communication Interfaces
PTP
PTP
STEP 1: Connect DCH and DWS to PTP
Cu, 1GbE
Cu, 1GbE
Detector
Control
Hardware
Detector
Workstation
Combined link
PTP + P-CND
is possible but not recommended.
PTP on DWS is not strictly necessary but strongly recommended.

## Slide 29: How to use it? :: Communication
Communication Interfaces
PTP
PTP
STEP 2: Connect maintenance network(s) of DWS
Management
Cu, 1GbE
Cu, 1GbE
Cu
Detector
Control
Hardware
Detector
Workstation
Follow ESO networks standards.

## Slide 30: How to use it? :: Communication
Communication Interfaces
Detector
Control
Hardware
Detector
Workstation
PTP
PTP
STEP 3: Connect both DCH and DWS to P-CND
Management
Cu, 1GbE
Cu, 1GbE
Cu
P-CND
Cu, 1GbE
/
Opt. 1GbE
DAC, 10GbE
DWS uplink must be high-bandwidth since it includes image data uplink.
Switches possible on platform / in computer room depending on instrument needs.
No specific requirements on P-CND switches.

## Slide 31: How to use it? :: Communication
Communication Interfaces
Detector
Control
Hardware
Detector
Workstation
PTP
PTP
STEP 4: Connect image data link from DCH to DWS through P-RT
Management
Cu, 1GbE
Cu, 1GbE
Cu
P-CND
Cu, 1GbE
/
Opt. 1GbE
DAC, 10GbE
P-RT
Opt. 10GbE
Most basic case:
Direct fibre connection from DCH to DWS
Preferred: add RT capable switches in computer room.

## Slide 32: Synchronization

## Slide 33: How to use it? :: Synchronization
White Rabbit
  White rabbit is a synchronization protocol developed at CERN
    https://white-rabbit.web.cern.ch/
    Sub-nanosecond accuracy and picoseconds precision of synchronization
    Typical distances of 10 km between network elements
  Now part of the PTP standard IEEE1588-2019 as “High Accuracy Profile”
  Advantages over coaxial TTL synchronization:
    No matched cables
    No practical limit to node distance
    No grounding issues

## Slide 34: How to use it? :: Synchronization
White Rabbit
CREOTECH (PL)
SAFRAN (FR)
GMV (DE)
The white rabbit switch is open-source hardware.
https://ohwr.org/project/wr-switch-hw-v4
Produced by multiple manufacturers.

## Slide 35: How to use it? :: Synchronization
White Rabbit
  Ideally:
    White Rabbit available as a telescope-wide service, like PTP
  Instruments need to create a local (to the platform) White Rabbit network.
    All NGCII Detector Control Hardware (DCH) units connect to a White Rabbit switch
    WR switch in free-running mode
    No uplink to other equipment
    Dedicated fibres! (1000BASE-BX10, Single LC Type)
This is (currently) not the case!
Data Classification: PUBLIC, ESO-598352 v.1

## Slide 36: Shutters

## Slide 37: How to use it? :: Shutters
Legacy shutter control interface of FIERA/NGC
NGCII/OPT
Software
NGCII/CORE
Software
NGCII
Firmware
PLC
Shutter
Instrument WS / Detector WS
Detector Control HW
TTL Interface ESO-044439
Legacy shutter control is considered obsolete for
NGCII in VLT and ELT

## Slide 38: How to use it? :: Shutters
New shutter control concept
NGCII/OPT
Software
NGCII/CORE
Software
NGCII
Firmware
PLC
Shutter
Instrument WS / Detector WS
Detector Control HW
Ethernet based protocol
Instrument
Control
Network
PTP

## Slide 39: Detector Interface

## Slide 40: How to use it? :: Detector Interface
Detector Interface
  One to six EN4165 connectors on the rear panel
    Mandatory locking. Requires 3mm hex key for mounting and unmounting
    Mechanical keying. No way to mis-connect in a way that damages the detector
The METIS project is working on a right-angle connector variant!

## Slide 41: Power

## Slide 42: How to use it? :: Power
Power
  Default configuration: 230V AC, C13 connector
  Power consumption in ICD ESO-566603, Table 4
    Typical power for dimensioning of input supply. (Apply generous margin!)
    Maximum power will not be exceeded by NGCII (over current protection will trip before)
  −48V DC power input available as an option.
    Special use-cases, consult ESO.
    External AC/DC converter required.

## Slide 43: Cooling

## Slide 44: How to use it? :: Cooling
Cooling
Front Side
Fan Tray
METIS vibration tests show:
Exported vibration can be kept below 0.2NRMS with reasonable effort put into cabinet design.
Vibration frequency is 108Hz to 164Hz depending on speed setting. → above most critical range for ELT Nasmyth for normal operation.

## Slide 45: When can I have one?

## Slide 46: When can I have one?
It depends
For delivery dates of your particular system please contact your ESO detector system engineer.

## Slide 47: When can I have one?
Grade of Hardware
Series
All functions performing to specifications including environmental tests.
Final production version.
Delivery to Europe and Chile.
Pre-Series
All functions performing to specifications.
No significant reworks.
May not yet have passed environmental testing.
Delivery to Europe possible.
Delivery to Chile after individual approval.
Prototype
Core functions performing to specifications.
Suitable for testing detectors.
May be missing some secondary functions.
May have significant reworks.
Delivery outside of ESO only under exceptional circumstances.
1-2 systems
~10 systems
∞ systems

## Slide 48: When can I have one?
It depends
Firmware Development
Concept
Prototyping
Integration
Testing
Pre-Series
Production
Series
Concept
Prototyping
Integration
Testing
Pre-Series
Series
HAWAII-xRG
SAPHIRA
CCD
GEOSNAP

## Slide 49: When can I have one?
Environmental Testing
  Ensure that NGCII will pass instrument level testing during integration.
  Testing three configurations covering all modules:
    SAPHIRA (also representative of HAWAII-xRG)
    CCD231 (also representative of other CCDs)
    GEOSNAP
Thermal Testing:
High Temperature Operation
Low Temperature Operation
Cold Start
EMC Pre-Compliance:
ESD
Radiated
Emission
Radiated
Immunity
Conducted
Immunity
(Conducted
Emission)
EN 300 019-1-3 Class 3.1

## Slide 50: When can I have one?
It depends
For delivery dates of your particular system please contact your ESO detector system engineer.

## Slide 52
@ESOAstronomy
@esoastronomy
@ESO
european-southern-observatory
@ESOobservatory
Mathias Richerzhagen
Matthias Seidel
Alexander Rüde
Omar Sqalli